## Commentary
As S21815.xls
Start with S21580a
Update history page for easier plotting.
Copy 'Universal Bridge' from most recent S21814 for RLC standards.
Carefully redefine critical block names.
Alter label structure on Results page, could try to use vlookup to find range zeros on UniversalBridge sheet?
Replace uncertainty sheet with the uncertainty sheet from S21814
Alter treatment of range 7Y and 6Y by applying corrections as at bottom of Results sheet
Modify Report to pick up corrected results
As S21942.xlsx
Start with S21815.xls
Enter Ubridge calibration data of 7/7/14
Add graph history of UB cal constants, just for interest.
Delays in calibrating the UB (and repairing the permutable) are the reason for having the standards so long, and hence measuring them twice.
Range 7 results remain a problem for loss angle measurements, see 'Range7' worksheet.
Therefore for reporting loss angle for the 1, 10 and 100 pF capacitors, use AH2700 results with UB uncertainties (for CMC compliance).
As there is doubt about the 1 June loss angle measurements, only the 12 August results are reported
As S21959
Start with S21942b.xls as an up to date UB calculator for capacitance and loss angle
Reference S21745_v3.xls for structure of previous calibration of the Olex capacitance box
Modify labels in results, labelling box with dials A to E (slightly different from S21745)
Update Zeros table in 'Universal Bridge' to pick up correct zeros from 'Results' page
Note that E0 and D0 will likely have had a dial movement and so would not be identical
Use trick of labelling with a slightly different frequency!!
Build report page afresh, copying format from S21745.
New value for UB1000 inserted in 'Unversal Bridge' sheet. Result of likely step change, see p.33 of UBIII
Set capacitance temperature coefficient uncertainty as 70 ppm one sigma, based on manufacture's claim that coefficient is -140 ppm.
Use 'connection definition' uncertainty as in S21745.
Formula for series resistance mapping as conductance brought into uncertainty from S21745
AH2700Results carries the AH2700 measurements of the 100 pF/step and 1pF/div dial.
Report page now selectively uses AH2700 results for phase angle for 100 pF dial down.
Update RoomLogger sheet from Planck1 logging system.
GRhistory page now set up to display history graph for each setting by entering the dial label in the highlighted box.
Conditions page just summarises temp and humidity.
As S21960
Start with S1959final (for capacitance) as most up to date UB corrections etc.
Look to S21784checkedTJS for report formatting for the Muirhead inductors.
Results page has inductance data entered with suitable labels for processing.
Universal Bridge page points to the relevant zeros in the Results page for these decade inductors.
Insert Lreport page based on S21784 report format.
Change uncertainty to point to ub_uncertainty block.
Use the new labels consistent with the Results page
Add Lhistory page to capture trends
Uncertainty, row 61 effect of series resistance on loss angle of a capacitor
Enter copper temperature coefficient in the parameters block
Did not add 3 month change as an uncertainty (c.f.S21784). Assume client will monitor history/stability.
Add CMC for inductors block at far right of Uncertainty page.
Capacitor worksheets grouped away from inductor worksheets.
As S1997
Transition to python analysis
Note that the item names 'coax zero' and 'coax open' are reserved names.
There should be a zero/open for each range and nominal frequency that is used.
The zero or open dial readings are calculated as an immittance to be subtracted from the measured immittances.
In some cases there will be alternative zeros specific to the UUT that are subtracted instead.
An essentially standalone python calculation takes data from 'pyUBreadings' to 'pyUBresults' from which a report is made in 'pyReport'.
py4284readings' contains results from the HP4284 as a cross-check, but not contributing directly to the report.
At the same time the data has also been copied into 'excelResults' and this has been processed, with minor updates, to give the report in 'Lreport'.
This uses UB constants and calculation in the 'Universal Bridge' sheet, with uncertainties calculated in the 'Uncertainty' sheet.
As S22012
For the Babcock Muirhead capacitors.
Post processing to cope with loss angle calculations and a 0.1 pF definition uncertainty.
This has been merged with S21986 for the cross-check with previous Excel calculations (see 'ExcelCommentary').
As S22013
For the Babcock HP16380 sets.
Adjust Excel uncertainty for 16380 temperature coefficients and to remove series resistance effect.
Minor change to Python results to mix in AH2700A loss angle results.
Python table points to pyUbesults as in S21997, i.e. no need for a special Python post processing.
HP history updated.
Higher python loss angle uncertainty due to including uncertainty from the zero, whereas Excel just subtracts the zero reading with no uncertainty.
Temperature in pyUBresults forced to 20. This avoids Python trying to correct the roughly 21 degree results back to 20, but leaves the correct uncertainty.

## pyUBreadings
Template for capacitors
Capacitors | Babcock HP16380 |  |  | Client | Babcock |  |  |  |  | Enter, say, 5 ppm as 5e-6
Date | 2017-11-20 |  |  | Job number | 96330108 |  |  |  |  | 1 sigma uncertainties from specifications
|  |  |  | S Number | S22013
|  |  | Data from lab book |  |  |  | Data from specifications and other sources
Item | Nominal Frequency | Range | xdial | rdial | Frequency | Temperature | u | df | Temp coeff C | u | df | Temp coeff R | u | df | Other | u | df
| Hz |  |  |  | Hz | °C
1 microF | 1.6k | 3Y | 999846 | 1716 | 1592.03 | 20 | 0.25 | 10 | 7.5e-06 | 2.75e-05 | 10 | 0 | 0 | 10
coax zero | 1.6k | 3Y | 0 | -2 | 1592.02 | 20 | 0.25 | 10 | 0 | 0 | 10 | 0 | 0 | 10
coax zero | 1.6k | 4Y | 0 | -1 | 1591.73 | 20 | 0.25 | 10 | 0 | 0 | 10 | 0 | 0 | 10
0.1 microF | 1.6k | 4Y | 999964 | 412 | 1591.65 | 20 | 0.25 | 10 | 7.5e-06 | 2.75e-05 | 10 | 0 | 0 | 10
0.01 microF | 1.6k | 5Y | 1000024 | 498 | 1591.58 | 20 | 0.25 | 10 | 7.5e-06 | 2.75e-05 | 10 | 0 | 0 | 10
coax zero | 1.6k | 5Y | -1 | 2 | 1591.52 | 20 | 0.25 | 10 | 0 | 0 | 10 | 0 | 0 | 10
coax zero | 1.6k | 6Y | 0 | 14 | 1591.77 | 20 | 0.25 | 10 | 0 | 0 | 10 | 0 | 0 | 10
1000 pF | 1.6k | 6Y | 999731 | -187 | 1591.75 | 20 | 0.25 | 10 | 3e-05 | 9e-06 | 10 | 0 | 0 | 10
GR1000A | 1.6k | 6Y | 999890 | -150 | 1591.67 | 20 | 0.25 | 10 | 0 | 0 | 10 | 0 | 0 | 10
GR1000B | 1.6k | 6Y | 999876 | -146 | 1591.6 | 20 | 0.25 | 10 | 0 | 0 | 10 | 0 | 0 | 10
AH11C | 1.6k | 6Y | 99985 | 104 | 1591.5 | 20 | 0.25 | 10 | 0 | 0 | 10 | 0 | 0 | 10
100 pF | 1.6k | 6Y | 99990 | -15 | 1591.42 | 20 | 0.25 | 10 | 3e-05 | 9e-06 | 10 | 0 | 0 | 10
100 pF | 1.6k | 7Y | 999946 | 1996 | 1591.32 | 20 | 0.25 | 10 | 3e-05 | 9e-06 | 10 | 0 | 0 | 10
AH11C | 1.6k | 7Y | 999901 | 2930 | 1591.19 | 20 | 0.25 | 10 | 0 | 0 | 10 | 0 | 0 | 10
coax zero | 1.6k | 7Y | 6 | 19 | 1591.12 | 20 | 0.25 | 10 | 0 | 0 | 10 | 0 | 0 | 10
AH11A | 1.6k | 7Y | 99995 | 201 | 1591.82 | 20 | 0.25 | 10 | 0 | 0 | 10 | 0 | 0 | 10
10 pF | 1.6k | 7Y | 99987 | 210 | 1591.77 | 20 | 0.25 | 10 | 3e-05 | 9e-06 | 10 | 0 | 0 | 10
1 pF | 1.6k | 7Y | 10004 | 46 | 1591.72 | 20 | 0.25 | 10 | 4e-05 | 1.2e-05 | 10 | 0 | 0 | 10

## pyUBresults
Item | Label | Range | xDial | rDial | frequency | temperature | utemp | dftemp | alphax | ualphax | dfalphax | alphar | ualphar | dfalphr | x | ux | kx | r | ur | kr | cmcx | cmcr | ux/cmcx | ur/cmcr | label | multiplier for C | C |  |  | S |  |  | tand(mrad) | u tan delta | k | AH2700A tand
1 microF | 1.6k | 3Y | 999846 | 1716 | 1592.03 | 20 | 0.25 | 10 | 7.5e-06 | 2.75e-05 | 10 | 0 | 0 | 10 | 9.9992e-07 | 4.5901e-11 | 2.0239 | 1.7179e-06 | 3.7844e-07 | 2.2184 | 2.1998e-11 | 2.7006e-07 | 2.0865 | 1.4013 | 1 microF1.6k | 1000000 | 0.99992 | 4.5901e-05 | 2.0239 | 1.7179e-06 | 3.7844e-07 | 2.2184 | 0.17175 | 0.037836 | 2.2184
coax zero | 1.6k | 3Y | 0 | -2 | 1592.02 | 20 | 0.25 | 10 | 0 | 0 | 10 | 0 | 0 | 10 | 0 | 0 | 0 | 0 | 0 | 0 | 1.2563e-16 | 0 | 0 | #DIV/0! | coax zero1.6k | 1000000 | 0 | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0
coax zero | 1.6k | 4Y | 0 | -1 | 1591.73 | 20 | 0.25 | 10 | 0 | 0 | 10 | 0 | 0 | 10 | 0 | 0 | 0 | 0 | 0 | 0 | 1.2565e-16 | 0 | 0 | #DIV/0! | coax zero1.6k | 1000000 | 0 | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0
0.1 microF | 1.6k | 4Y | 999964 | 412 | 1591.65 | 20 | 0.25 | 10 | 7.5e-06 | 2.75e-05 | 10 | 0 | 0 | 10 | 1e-07 | 4.5906e-12 | 2.0239 | 4.1298e-08 | 3.7831e-08 | 2.2184 | 2.2002e-12 | 2.7005e-08 | 2.0864 | 1.4009 | 0.1 microF1.6k | 1000000 | 0.1 | 4.5906e-06 | 2.0239 | 4.1298e-08 | 3.7831e-08 | 2.2184 | 0.041293 | 0.037827 | 2.2184
0.01 microF | 1.6k | 5Y | 1000024 | 498 | 1591.58 | 20 | 0.25 | 10 | 7.5e-06 | 2.75e-05 | 10 | 0 | 0 | 10 | 1.0001e-08 | 4.5909e-13 | 2.0239 | 4.9597e-09 | 3.783e-09 | 2.2184 | 2.2015e-13 | 2.703e-09 | 2.0854 | 1.3995 | 0.01 microF1.6k | 1000000 | 0.010001 | 4.5909e-07 | 2.0239 | 4.9597e-09 | 3.783e-09 | 2.2184 | 0.049591 | 0.037825 | 2.2184
coax zero | 1.6k | 5Y | -1 | 2 | 1591.52 | 20 | 0.25 | 10 | 0 | 0 | 10 | 0 | 0 | 10 | 0 | 0 | 0 | 0 | 0 | 0 | 1.2545e-16 | 2.6972e-12 | 0 | 0 | coax zero1.6k | 1000000 | 0 | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0
coax zero | 1.6k | 6Y | 0 | 14 | 1591.77 | 20 | 0.25 | 10 | 0 | 0 | 10 | 0 | 0 | 10 | 0 | 0 | 0 | 0 | 0 | 0 | 1.2565e-16 | 0 | 0 | #DIV/0! | coax zero1.6k | 1000000000000 | 0 | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0
1000 pF | 1.6k | 6Y | 999731 | -187 | 1591.75 | 20 | 0.25 | 10 | 3e-05 | 9e-06 | 10 | 0 | 0 | 10 | 9.9987e-10 | 4.8907e-14 | 2.0416 | -2.01e-10 | 3.7829e-10 | 2.2184 | 2.2123e-14 | 2.727e-10 | 2.2107 | 1.3872 | 1000 pF1.6k | 1000000000000 | 999.8689 | 0.048907 | 2.0416 | -2.01e-10 | 3.7829e-10 | 2.2184 | -0.0201 | 0.037829 | 2.2184
GR1000A | 1.6k | 6Y | 999890 | -150 | 1591.67 | 20 | 0.25 | 10 | 0 | 0 | 10 | 0 | 0 | 10 | 1e-09 | 3.3334e-14 | 2.1262 | -1.64e-10 | 3.7831e-10 | 2.2184 | 2.2126e-14 | 2.7273e-10 | 1.5065 | 1.3871 | GR1000A1.6k | 1000000000000 | 1000.0279 | 0.033334 | 2.1262 | -1.64e-10 | 3.7831e-10 | 2.2184 | -0.016398 | 0.037828 | 2.2184 | 0.0087458
GR1000B | 1.6k | 6Y | 999876 | -146 | 1591.6 | 20 | 0.25 | 10 | 0 | 0 | 10 | 0 | 0 | 10 | 1e-09 | 3.3333e-14 | 2.1262 | -1.6e-10 | 3.7828e-10 | 2.2184 | 2.2126e-14 | 2.7271e-10 | 1.5065 | 1.3871 | GR1000B1.6k | 1000000000000 | 1000.0139 | 0.033333 | 2.1262 | -1.6e-10 | 3.7828e-10 | 2.2184 | -0.015999 | 0.037826 | 2.2184
AH11C | 1.6k | 6Y | 99985 | 104 | 1591.5 | 20 | 0.25 | 10 | 0 | 0 | 10 | 0 | 0 | 10 | 9.9999e-11 | 5.3758e-15 | 2.0059 | 9e-11 | 5.5765e-11 | 2.0187 | 2.3256e-15 | 2.9699e-11 | 2.3115 | 1.8777 | AH11C1.6k | 1000000000000 | 99.9988 | 0.0053758 | 2.0059 | 9e-11 | 5.5765e-11 | 2.0187 | 0.090004 | 0.055767 | 2.0187 | 0.0015881
100 pF | 1.6k | 6Y | 99990 | -15 | 1591.42 | 20 | 0.25 | 10 | 3e-05 | 9e-06 | 10 | 0 | 0 | 10 | 1e-10 | 6.4601e-15 | 1.9969 | -2.9e-11 | 5.5764e-11 | 2.0187 | 2.3258e-15 | 2.9699e-11 | 2.7776 | 1.8777 |  | 1000000000000 | 100.0038 | 0.0064601 | 1.9969 | -2.9e-11 | 5.5764e-11 | 2.0187 | -0.029001 | 0.055767 | 2.0187
100 pF | 1.6k | 7Y | 999946 | 1996 | 1591.32 | 20 | 0.25 | 10 | 3e-05 | 9e-06 | 10 | 0 | 0 | 10 | 1e-10 | 4.8914e-15 | 2.0416 | 1.9769e-10 | 3.7814e-11 | 2.2184 | 2.3257e-15 | 2.9696e-11 | 2.1032 | 1.2734 | 100 pF1.6k | 1000000000000 | 100.0003 | 0.0048914 | 2.0416 | 1.9769e-10 | 3.7814e-11 | 2.2184 | 0.19771 | 0.037819 | 2.2184 | 0.00043
AH11C | 1.6k | 7Y | 999901 | 2930 | 1591.19 | 20 | 0.25 | 10 | 0 | 0 | 10 | 0 | 0 | 10 | 9.9996e-11 | 3.3331e-15 | 2.1262 | 2.9108e-10 | 3.7806e-11 | 2.2184 | 2.3256e-15 | 2.9692e-11 | 1.4332 | 1.2733 | AH11C1.6k | 1000000000000 | 99.9958 | 0.0033331 | 2.1262 | 2.9108e-10 | 3.7806e-11 | 2.2184 | 0.29116 | 0.037816 | 2.2184 | 0.0015881
coax zero | 1.6k | 7Y | 6 | 19 | 1591.12 | 20 | 0.25 | 10 | 0 | 0 | 10 | 0 | 0 | 10 | 0 | 0 | 0 | 0 | 0 | 0 | 1.2571e-16 | 2.6994e-12 | 0 | 0 | coax zero1.6k | 1000000000000 | 0 | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0
AH11A | 1.6k | 7Y | 99995 | 201 | 1591.82 | 20 | 0.25 | 10 | 0 | 0 | 10 | 0 | 0 | 10 | 9.9995e-12 | 5.3756e-16 | 2.0059 | 1.8199e-11 | 5.5772e-12 | 2.0188 | 3.4565e-16 | 5.401e-12 | 1.5552 | 1.0326 | AH11A1.6k | 1000000000000 | 9.9995 | 0.00053756 | 2.0059 | 1.8199e-11 | 5.5772e-12 | 2.0188 | 0.18196 | 0.055765 | 2.0188 | -0.00078503
10 pF | 1.6k | 7Y | 99987 | 210 | 1591.77 | 20 | 0.25 | 10 | 3e-05 | 9e-06 | 10 | 0 | 0 | 10 | 9.9987e-12 | 6.4596e-16 | 1.9969 | 1.9099e-11 | 5.5769e-12 | 2.0187 | 3.4563e-16 | 5.4006e-12 | 1.8689 | 1.0327 | 10 pF1.6k | 1000000000000 | 9.9987 | 0.00064596 | 1.9969 | 1.9099e-11 | 5.5769e-12 | 2.0187 | 0.19098 | 0.055769 | 2.0187 | -0.0019472
1 pF | 1.6k | 7Y | 10004 | 46 | 1591.72 | 20 | 0.25 | 10 | 4e-05 | 1.2e-05 | 10 | 0 | 0 | 10 | 9.9986e-13 | 4.4518e-16 | 2.0206 | 2.6998e-12 | 4.4299e-12 | 2.022 | 1.4766e-16 | 2.9704e-12 | 3.0149 | 1.4913 | 1 pF1.6k | 1000000000000 | 0.99986 | 0.00044518 | 2.0206 | 2.6998e-12 | 4.4299e-12 | 2.022 | 0.26999 | 0.443 | 2.022 | -0.01404
correct for 0 mrad on AH
[15 empty] | 6Y | 1.0001e-15 | 1.0001 | 137.8985
[15 empty] | 6Y | 1.0001e-15 | 1.0001 | 137.8985
[15 empty] | 6Y | 1.0001e-15 | 1.0001 | 137.8985
[15 empty] | 6Y | 1.0001e-15 | 1.0001 | 137.8985
[15 empty] | 6Y | 1.0001e-15 | 1.0001 | 137.8985

## AH27000A
Measurements were also made with the AH2700 at 1591.5 Hz
page 70 of AIVII |  |  |  |  | 1591.5
2017-11-20 | Item | Capacitance | Conductance |  | tand
|  | pF | nS |  | mrad
| GR1000A | 1000.0308 | 0.08746 |  | 0.0087458
| AH11A | 10 | -7.85e-05 |  | -0.00078503
| AH11C | 99.9998 | 0.001588 |  | 0.0015881
| 1 pF | 0.99994 | -0.0001404 | 20.75 | -0.01404
| 10 pF | 9.9992 | -0.0001947 | 20.75 | -0.0019472
| 100 pF | 100.0042 | 0.00043 | 20.75 | 0.00043
| 1000 pF | 999.8629 | 0.01197 | 20.77 | 0.0011972

## pyReport
| Python | [7 empty] | Excel | [6 empty] | Differences
|  | Nominal Value (pF) | Capacitance (pF) | Expanded Uncertainty (pF) | k | Loss Angle (mrad) | Expanded Uncertainty (mrad) | k | Nominal Value (pF) | Capacitance (pF) | Expanded Uncertainty (pF) | k | Loss Angle (mrad) | Expanded Uncertainty (mrad) | k | Nominal Value (pF) | Capacitance (pF) | Expanded Uncertainty (pF) | k | Loss Angle (mrad) | Expanded Uncertainty (mrad) | k
1000 pF1.6k |  | 1000 | 999.8689 | 0.048907 | 2.0416 | 0 | 0.037829 | 2.2184 | 1000 | 999.8688 | 0.050273 | 2.1371 | -0.020103 | 0.034531 | 2.15 | 0 | 3.2259e-05 | -0.0013653 | -0.095558 | 0.020103 | 0.0032982 | 0.068324
100 pF1.6k |  | 100 | 100.0003 | 0.0048914 | 2.0416 | 0.00043 | 0.037819 | 2.2184 | 100 | 100.0003 | 0.0050283 | 2.1371 | 0.19769 | 0.034534 | 2.15 | 0 | 3.7201e-06 | -0.00013694 | -0.095559 | -0.19726 | 0.0032854 | 0.068313
10 pF1.6k |  | 10 | 9.9987 | 0.00064596 | 1.9969 | -0.0019472 | 0.055769 | 2.0187 | 10 | 9.9987 | 0.00059597 | 2.0955 | 0.19102 | 0.044021 | 2.0613 | 0 | 3.7196e-07 | 4.9982e-05 | -0.098651 | -0.19296 | 0.011748 | -0.042532
1 pF1.6k |  | 1 | 0.99986 | 0.00044518 | 2.0206 | -0.01404 | 0.443 | 2.022 | 1 | 0.99986 | 0.00041578 | 2.5532 | 0.27003 | 0.31071 | 2.1827 | 0 | 3.7196e-08 | 2.9402e-05 | -0.53259 | -0.28407 | 0.13229 | -0.16078
|  | Nominal Value (µF) | Capacitance (µF) | Expanded Uncertainty (µF) | k | Loss Angle (mrad) | Expanded Uncertainty (mrad) | k | Nominal Value (µF) | Capacitance (µF) | Expanded Uncertainty (µF) | k | Loss Angle (mrad) | Expanded Uncertainty (mrad) | k | Nominal Value (µF) | Capacitance (µF) | Expanded Uncertainty (µF) | k | Loss Angle (mrad) | Expanded Uncertainty (mrad) | k
1 microF1.6k |  | 1 | 0.99992 | 4.5901e-05 | 2.0239 | 0.17175 | 0.037836 | 2.2184 | 1 | 0.99992 | 4.2879e-05 | 2.2941 | 0.17181 | 0.034533 | 2.15 | 0 | 3.9404e-08 | 3.0212e-06 | -0.2702 | -5.72e-05 | 0.0033026 | 0.068333
0.1 microF1.6k |  | 0.1 | 0.1 | 4.5906e-06 | 2.0239 | 0.041293 | 0.037827 | 2.2184 | 0.1 | 0.1 | 4.2884e-06 | 2.2941 | 0.041297 | 0.034533 | 2.15 | 0 | 3.9409e-09 | 3.0216e-07 | -0.2702 | -3.8924e-06 | 0.0032937 | 0.068324
0.01 microF1.6k |  | 0.01 | 0.010001 | 4.5909e-07 | 2.0239 | 0.049591 | 0.037825 | 2.2184 | 0.01 | 0.010001 | 4.2887e-07 | 2.2941 | 0.049594 | 0.034533 | 2.15 | 0 | 3.9411e-10 | 3.0218e-08 | -0.2702 | -2.4934e-06 | 0.0032921 | 0.068322

## HP history
HISTORY |  |  | Checked by: | ACC
Instructions
Select rows 24:33 and carry out "Insert Rows". Copy cells in green block A14:N22 and "Paste" to cells A24:N32. Copy cells in blue block and "Paste Special, Values" to green block A14:N22.
2017-11-10 | S22013
HP 16387A 1 µF | [7 empty] | 0.99992 | 4.5901e-05 | 2.0239 | 0.17175 | 0.037836 | 2.2184
HP 16386A 0.1 µF | [7 empty] | 0.1 | 4.5906e-06 | 2.0239 | 0.041293 | 0.037827 | 2.2184
HP 16385A 0.01 µF | [7 empty] | 0.010001 | 4.5909e-07 | 2.0239 | 0.049591 | 0.037825 | 2.2184
HP 16384A 1000 pF | [7 empty] | 999.8689 | 0.048907 | 2.0416 | 0 | 0.037829 | 2.2184
HP 16383A 100 pF | [7 empty] | 100.0003 | 0.0048914 | 2.0416 | 0.00043 | 0.037819 | 2.2184
HP 16382A 10 pF | [7 empty] | 9.9987 | 0.00064596 | 1.9969 | -0.0019472 | 0.055769 | 2.0187
HP 16381A 1 pF | [7 empty] | 0.99986 | 0.00044518 | 2.0206 | -0.01404 | 0.443 | 2.022
2014-08-12 | S21942 | [5 empty] | Nominal Value | Capacitance | Uncertainty | Coverage | Loss Angle | Uncertainty | Coverage
|  | Range | C | G |  |  | (mF) or (pF) | (mF) or (pF) | (mF) or (pF) | Factor | (milliradians) | (milliradians) | Factor |  | change | unc ppm |  |  |  |  | Lookup Label | Previous | Current | Days |  |  |  |  | 3 month change, ppm
HP 16387A 1 µF |  | 3Y | 999801 | 1696 |  |  |  | 0.9999 | 4.2891e-05 | 2.2948 | 0.17091 | 0.034534 | 2.15 |  | 16.3006 | 42.895 |  |  |  |  | A1 | 2005-02-22 | 2014-08-12 | 3410 |  |  |  |  | 0.43022
HP 16386A 0.1 µF |  | 4Y | 999928 | 394 |  |  |  | 0.1 | 4.2896e-06 | 2.2948 | 0.041297 | 0.034534 | 2.15 |  | 6.2985 | 42.895 |  |  |  |  | B10 | 2005-02-22 | 2014-08-12 | 3410 |  |  |  |  | 0.16624
HP 16385A 0.01 µF |  | 5Y | 999986 | 496 |  |  |  | 0.010001 | 4.2899e-07 | 2.2948 | 0.050094 | 0.034533 | 2.15 |  | 17.298 | 42.8949 |  |  |  |  | C10 | 2005-02-22 | 2014-08-12 | 3410 |  |  |  |  | 0.45655
HP 16384A 1000 pF |  | 6Y | 999660 | -205 |  |  |  | 999.8274 | 0.050849 | 2.1617 | 0.00028984 | 0.034531 | 2.15 |  | -7.2241 | 50.8575 |  |  |  |  | D1 | 2005-02-22 | 2014-08-12 | 3410 |  |  |  |  | -0.19067
HP 16383A 100 pF |  | 7Y | 999865 | 2540 |  |  |  | 100.0022 | 0.0050861 | 2.1617 | 0.00040501 | 0.034534 | 2.15 |  | 111.3135 | 50.86 |  |  |  |  | E1 | 2005-02-22 | 2014-08-12 | 3410 |  |  |  |  | 2.9379
HP 16382A 10 pF |  | 7Y | 99969 | 260 |  |  |  | 9.9981 | 0.00059909 | 2.1066 | -0.0015382 | 0.044023 | 2.0613 |  | -63.7133 | 59.9205 |  |  |  |  | E2 | 2005-02-22 | 2014-08-12 | 3410 |  |  |  |  | -1.6816
HP 16381A 1 pF |  | 7Y | 10004 | 40 |  |  |  | 0.99989 | 0.00041579 | 2.5532 | -0.009751 | 0.31071 | 2.1827 |  | 736.4257 | 415.8365 |  |  |  |  | E3 | 2005-02-22 | 2014-08-12 | 3410 |  |  |  |  | 19.4365
2011-05-17 | S21815 | [5 empty] | Nominal Value | Capacitance | Uncertainty | Coverage | Loss Angle | Uncertainty | Coverage
|  | Range | C | G |  |  | (mF) or (pF) | (mF) or (pF) | (mF) or (pF) | Factor | (milliradians) | (milliradians) | Factor
HP 16387A 1 µF |  | 3Y | 999831 | 1710 |  |  |  | 0.9999 | 4.2891e-05 | 2.1617 | 0.17111 | 0.034534 | 2.15
HP 16386A 0.1 µF |  | 4Y | 999959 | 397 |  |  |  | 0.1 | 4.2896e-06 | 2.1617 | 0.039798 | 0.034534 | 2.15
HP 16385A 0.01 µF |  | 5Y | 1000016 | 505 |  |  |  | 0.010001 | 4.2899e-07 | 2.1066 | 0.050195 | 0.034533 | 2.15
HP 16384A 1000 pF |  | 6Y | 999702 | -190 |  |  |  | 999.8369 | 0.050849 | 2.1617 | -0.0055039 | 0.034531 | 2.15
HP 16383A 100 pF |  | 7Y | 999888 | 2100 |  |  |  | 100.0007 | 0.0050859 | 2.1617 | 0.0055476 | 0.034534 | 2.15
HP 16382A 10 pF |  | 7Y | 99974 | 213 |  |  |  | 9.9979 | 0.00059907 | 2.1066 | -0.00791 | 0.044023 | 2.0613
HP 16381A 1 pF |  | 7Y | 10005 | 36 |  |  |  | 1 | 0.00041579 | 2.5532 | 0.0040466 | 0.31068 | 2.1827
2005-02-22 | S21580 | [5 empty] | Nominal Value | Capacitance | Uncertainty | Coverage | Loss Angle | Uncertainty | Coverage
|  | Range | C | G |  |  | (mF) or (pF) | (mF) or (pF) | (mF) or (pF) | Factor | (milliradians) | (milliradians) | Factor
HP 16387A 1 µF |  | 3Y | 999820 | 1653 |  |  |  | 0.99989 | 0.00011135 | 2.1283 | 0.16474 | 0.034393 | 2.15
HP 16386A 0.1 µF |  | 4Y | 999958 | 335 |  |  |  | 0.1 | 1.1136e-05 | 2.1262 | 0.033462 | 0.034393 | 2.15
HP 16385A 0.01 µF |  | 5Y | 1000004 | 500 |  |  |  | 0.010001 | 1.1141e-06 | 2.1038 | 0.049394 | 0.034393 | 2.15
HP 16384A 1000 pF |  | 6Y | 999703 | -227 |  |  |  | 999.8347 | 0.11137 | 2.1283 | -0.013604 | 0.034391 | 2.15
HP 16383A 100 pF |  | 7Y | 999863 | 1834 |  |  |  | 99.991 | 0.011161 | 2.1262 | -0.009865 | 0.034394 | 2.15
HP 16382A 10 pF |  | 7Y | 99988 | 193 |  |  |  | 9.9987 | 0.0011449 | 2.1038 | -0.0046376 | 0.043842 | 2.0613
HP 16381A 1 pF |  | 7Y | 9997 | 42 |  |  |  | 0.99915 | 0.00035273 | 2.0761 | 0.096162 | 0.30966 | 2.1828
2002-07-22 | S21385 | [5 empty] | Nominal Value | Capacitance | Uncertainty | Coverage | Loss Angle | Uncertainty | Coverage
Component |  | Range | Capacitance | Conductance |  |  | (mF) or (pF) | (mF) or (pF) | (mF) or (pF) | Factor | (milliradians) | (milliradians) | Factor
HP 16387A 1 µF |  | 3Y | 999800 | 1564 |  |  | 1 | 0.99989 | 0.00012347 | 2.4531 | 0.17381 | 0.039381 | 2.0872
HP 16386A 0.1 µF |  | 4Y | 999865 | 494 |  |  | 0.1 | 0.1 | 1.2347e-05 | 2.4545 | 0.060496 | 0.039381 | 2.0872
HP 16385A 0.01 µF |  | 4Y | 100000 | 54 |  |  | 0.01 | 0.010001 | 1.255e-06 | 2.3205 | 0.068991 | 0.03431 | 2.1402
HP 16384A 1000 pF |  | 6Y | 999725 | -96 |  |  | 1000 | 999.8515 | 0.074102 | 2.4353 | 0.0029977 | 0.03431 | 2.1402
HP 16383A 100 pF |  | 7Y | 999884 | 1530 |  |  | 100 | 99.9955 | 0.0074095 | 2.4406 | 0.0040066 | 0.03431 | 2.1402
HP 16382A 10 pF |  | 7Y | 100007 | 150 |  |  | 10 | 9.9993 | 0.00076244 | 2.09 | 0.019008 | 0.03431 | 2.1402
HP 16381A 1 pF |  | 7Y | 10027 | 24 |  |  | 1 | 0.99996 | 0.00049631 | 2.3634 | 0.026003 | 0.1071 | 2.1416
15/8/2000
Component |  | Range | C | G | True C |  | True G |  | tan _
HP 16387A 1 µF |  | 3Y | 999806 | 1734 | 0.99954 | µF | 0.0017336 | mS | 0.00017344
HP 16386A 0.1 µF |  | 4Y | 999945 | 600 | 100.0076 | nF | 0.059997 | µS | 5.9992e-05
HP 16385A 0.01 µF |  | 4Y | 100001 | 64 | 10.0014 | nF | 0.0063997 | µS | 6.3988e-05
HP 16384A 1000 pF |  | 6Y | 999733 | 14 | 999.4791 | pF | 0.013997 | nS | 1.4004e-06
HP 16383A 100 pF |  | 7Y | 999906 | 1080 | 103.4 | pF | 0.11178 | nS | 0.0001081
HP 16382A 10 pF |  | 7Y | 100012 | 110 | 10.3422 | pF | 0.011385 | nS | 0.00011008
HP 16381A 1 pF |  | 7Y | 10028 | 15 | 1.037 | pF | 0.0015525 | nS | 0.00014971
Results 27 November 1997
Component |  | Range | Capcitance | Conductance
HP 16387A 1 µF |  | 3Y | 999768 | 1756
HP 16386A 0.1 µF |  | 4Y | 999907 | 610
HP 16385A 0.01 µF |  | 5Y | 99998 | 56
HP 16384A 1000 pF |  | 6Y | 999756 | 98
HP 16383A 100 pF |  | 7Y | 999911 | 1100
HP 16382A 10 pF |  | 7Y | 100008 | 120
HP 16381A 1 pF |  | 7Y | 10028 | 10
Results 17 October 1995
Component |  | Range | Capcitance | Conductance
HP 16387A 1 µF |  | 3Y | 999775 | 1750
HP 16386A 0.1 µF |  | 4Y | 999904 | 590
HP 16385A 0.01 µF |  | 5Y | 1000254 | 700
HP 16384A 1000 pF |  | 6Y | 999770 | 69
HP 16383A 100 pF |  | 7Y | 1000019 | 1506
HP 16382A 10 pF |  | 7Y | 100020 | 150
HP 16381A 1 pF |  | 7Y | 10028 | 16
Results 5 August 1993
Component |  | Range | Capcitance | Conductance
HP 16387A 1 µF |  | 3Y | 999746 | 1813
HP 16386A 0.1 µF |  | 4Y | 999882 | 659
HP 16385A 0.01 µF |  | 5Y | 1000131 | 560
HP 16384A 1000 pF |  | 6Y | 999732 | 113
HP 16383A 100 pF |  | 7Y | 999893 | 1520
HP 16382A 10 pF |  | 7Y | 100014 | 150
HP 16381A 1 pF |  | 7Y | 10027 | 16
Date | Job | Report S No.
1993-08-05 |  | 20878
1995-10-17 | 71623 | 21066
1997-11-27 | 71750 | 21250
2000-08-15 | 73040 | 21385
2002-07-22 | 71250 | 21464
2005-02-22 | 7140032 | 21580
| 1993-08-04 |  |  |  | 1995-10-17 |  |  |  | 1997-11-27 |  |  |  | 2000-08-09 |  |  |  | 2002-07-22 |  |  |  | 2005-02-22 |  |  |  | 2011-05-17 |  |  |  | 2014-08-12 |  |  |  | 2017-11-20
| C | uC | delta | udelta | C | uC | delta | udelta | C | uC | delta | udelta | C | uC | delta | udelta | C | uC | delta | udelta | C | uC | delta | udelta | C | uC | delta | udelta | C | uC | delta | udelta | C | uC | delta | udelta
HP 16387A 1 µF | 0.99984 | 0.00019997 | 0.18 | 0.1 | 0.99986 | 0.00012 | 0.175 | 0.039 | 0.99985 | 0.00012 | 0.176 | 0.039 | 0.99989 | 0.00012 | 0.17 | 0.04 | 0.99988 | 0.00012 | 0.16 | 0.04 | 0.99989 | 0.00011135 | 0.16474 | 0.034393 | 0.9999 | 4.2891e-05 | 0.17111 | 0.034534 | 0.9999 | 4.2891e-05 | 0.17091 | 0.034534 | 0.99992 | 4.5901e-05 | 0.17175 | 0.037836
HP 16386A 0.1 µF | 0.099997 | 1.9999e-05 | 0.07 | 0.1 | 0.099998 | 1.2e-05 | 0.059 | 0.039 | 0.099999 | 1.2e-05 | 0.061 | 0.039 | 0.1 | 1.2e-05 | 0.06 | 0.04 | 0.1 | 1.2e-05 | 0.05 | 0.04 | 0.1 | 1.1136e-05 | 0.033462 | 0.034393 | 0.1 | 4.2896e-06 | 0.039798 | 0.034534 | 0.1 | 4.2896e-06 | 0.041297 | 0.034534 | 0.1 | 4.5906e-06 | 0.041293 | 0.037827
HP 16385A 0.01 µF | 0.01 | 2.0001e-06 | 0.08 | 0.1 | 0.010003 | 1.3e-06 | 0.064 | 0.034 | 0.010001 | 1.3e-06 | 0.059 | 0.034 | 0.010001 | 1.3e-06 | 0.07 | 0.04 | 0.010001 | 1.3e-06 | 0.06 | 0.03 | 0.010001 | 1.1141e-06 | 0.049394 | 0.034393 | 0.010001 | 4.2899e-07 | 0.050195 | 0.034533 | 0.010001 | 4.2899e-07 | 0.050094 | 0.034533 | 0.010001 | 4.5909e-07 | 0.049591 | 0.037825
HP 16384A 1000 pF | 999.87 | 0.19997 | 0 | 0.1 | 999.9 | 0.07 | 0 | 0.034 | 999.89 | 0.07 | 0 | 0.034 | 999.85 | 0.07 | 0 | 0.03 | 999.85 | 0.07 | 0.01 | 0.03 | 999.8347 | 0.11137 | -0.013604 | 0.034391 | 999.8369 | 0.050849 | -0.0055039 | 0.034531 | 999.8274 | 0.050849 | 0.00028984 | 0.034531 | 999.8689 | 0.048907 | 0 | 0.037829
HP 16383A 100 pF | 99.993 | 0.019999 | 0 | 0.1 | 100.005 | 0.008 | 0 | 0.034 | 99.994 | 0.008 | 0 | 0.034 | 99.995 | 0.007 | 0 | 0.03 | 99.992 | 0.007 | 0.07 | 0.03 | 99.991 | 0.011161 | -0.009865 | 0.034394 | 100.0007 | 0.0050859 | 0.0055476 | 0.034534 | 100.0022 | 0.0050861 | 0.00040501 | 0.034534 | 100.0003 | 0.0048914 | 0.00043 | 0.037819
HP 16382A 10 pF | 9.9989 | 0.0019998 | 0 | 0.1 | 9.9997 | 0.0008 | 0 | 0.034 | 9.9984 | 0.0008 | 0 | 0.034 | 9.9993 | 0.0008 | 0.02 | 0.03 | 9.9986 | 0.0008 | 0.06 | 0.03 | 9.9987 | 0.0011449 | -0.0046376 | 0.043842 | 9.9979 | 0.00059907 | -0.00791 | 0.044023 | 9.9981 | 0.00059909 | -0.0015382 | 0.044023 | 9.9987 | 0.00064596 | -0.0019472 | 0.055769
HP 16381A 1 pF | 0.99957 | 0.00019991 | 0.16 | 0.1 | 0.99996 | 0.00026 | 0.16 | 0.11 | 0.99986 | 0.0005 | 0.1 | 0.11 | 0.99996 | 0.0005 | 0.03 | 0.11 | 0.99986 | 0.0005 | 0.09 | 0.11 | 0.99915 | 0.00035273 | 0.096162 | 0.30966 | 1 | 0.00041579 | 0.0040466 | 0.31068 | 0.99989 | 0.00041579 | -0.009751 | 0.31071 | 0.99986 | 0.00044518 | -0.01404 | 0.443
HP 16387A 1 µF
Date | C | uC | delta | udelta
1993-08-05 | 0.99984 | 0.00019997 | 0.18 | 0.1
1995-10-17 | 0.99986 | 0.00012 | 0.175 | 0.039
1997-11-27 | 0.99985 | 0.00012 | 0.176 | 0.039
2000-08-15 | 0.99989 | 0.00012 | 0.17 | 0.04
2002-07-22 | 0.99988 | 0.00012 | 0.16 | 0.04
2005-02-22 | 0.99989 | 0.00011135 | 0.16474 | 0.034393
2011-05-17 | 0.9999 | 4.2891e-05 | 0.17111 | 0.034534
2014-08-12 | 0.9999 | 4.2891e-05 | 0.17091 | 0.034534
2017-11-20 | 0.99992 | 4.5901e-05 | 0.17175 | 0.037836
Note on X112: k.jones: set to zero in report for all four low value caps.

## ExcelCommentary
As S21815.xls
Start with S21580a
Update history page for easier plotting.
Copy 'Universal Bridge' from most recent S21814 for RLC standards.
Carefully redefine critical block names.
Alter label structure on Results page, could try to use vlookup to find range zeros on UniversalBridge sheet?
Replace uncertainty sheet with the uncertainty sheet from S21814
Alter treatment of range 7Y and 6Y by applying corrections as at bottom of Results sheet
Modify Report to pick up corrected results
As S21942.xlsx
Start with S21815.xls
Enter Ubridge calibration data of 7/7/14
Add graph history of UB cal constants, just for interest.
Delays in calibrating the UB (and repairing the permutable) are the reason for having the standards so long, and hence measuring them twice.
Range 7 results remain a problem for loss angle measurements, see 'Range7' worksheet.
Therefore for reporting loss angle for the 1, 10 and 100 pF capacitors, use AH2700 results with UB uncertainties (for CMC compliance).
As there is doubt about the 1 June loss angle measurements, only the 12 August results are reported
As S21959
Start with S21942b.xls as an up to date UB calculator for capacitance and loss angle
Reference S21745_v3.xls for structure of previous calibration of the Olex capacitance box
Modify labels in results, labelling box with dials A to E (slightly different from S21745)
Update Zeros table in 'Universal Bridge' to pick up correct zeros from 'Results' page
Note that E0 and D0 will likely have had a dial movement and so would not be identical
Use trick of labelling with a slightly different frequency!!
Build report page afresh, copying format from S21745.
New value for UB1000 inserted in 'Unversal Bridge' sheet. Result of likely step change, see p.33 of UBIII
Set capacitance temperature coefficient uncertainty as 70 ppm one sigma, based on manufacture's claim that coefficient is -140 ppm.
Use 'connection definition' uncertainty as in S21745.
Formula for series resistance mapping as conductance brought into uncertainty from S21745
AH2700Results carries the AH2700 measurements of the 100 pF/step and 1pF/div dial.
Report page now selectively uses AH2700 results for phase angle for 100 pF dial down.
Update RoomLogger sheet from Planck1 logging system.
GRhistory page now set up to display history graph for each setting by entering the dial label in the highlighted box.
Conditions page just summarises temp and humidity.
As S21986MuirheadCaps2016
Starting with S21959final
Note there is no need to invoke AH2700 for mesurements on Muirhead caps (500 pF to 1 microF)
Look to match notation of previous S21759_post_check.xls in 2010 where possible.
Modify (again) the labels for zero pickup in 'Universal Bridge' page. Now use 'differentzero' column to differentiate rather than previous fiddling with frequency.
New value for UB1000 inserted in 'Unversal Bridge' sheet. Result of likely step change, see p.239 of KJ lab diary 1
Formatting of result table modified to put all values in nF and to comply with decimal point alignment and minimalist use of lines
Review uncertainty specific to Muirhead specifications.
Connection definition changed to 0.1 pF/2, as in S21759 (but note that S21759 was still using % rather than counts).

## Room Conditions
2017-11-20 | 21.2 | 46.7 |  | RH correction | -2.128 | from Y:\Quality\TECHPROC\E058-Lab Air Cond\loggercalibration_2017.xls
2017-11-20 00:05 | 21.2 | 46.7 | Temperature | RH
2017-11-20 00:10 | 21.2 | 46.7 | 20.77 | 48.4039 | 46.2759
2017-11-20 00:15 | 21.2 | 46.7 | 0.091378 | 0.62736
2017-11-20 00:20 | 21.3 | 46.5
2017-11-20 00:25 | 21.2 | 46.5 | from GMH
2017-11-20 00:30 | 21.2 | 46.5
2017-11-20 00:35 | 21.2 | 46.4
2017-11-20 00:40 | 21.2 | 46.4
2017-11-20 00:45 | 21.2 | 46.3
2017-11-20 00:50 | 21.2 | 46.4
2017-11-20 00:55 | 21.2 | 46.3 |  |  |  |  | 1 microF | 1.6k | 3Y | 999846 | 1716 | 1592.03 | 20.86
2017-11-20 01:00 | 21.2 | 46.3 |  |  |  |  | coax zero | 1.6k | 3Y | 0 | -2 | 1592.02 | 20.85
2017-11-20 01:05 | 21.2 | 46.2 |  |  |  |  | coax zero | 1.6k | 4Y | 0 | -1 | 1591.73 | 20.85
2017-11-20 01:10 | 21.2 | 46.1 |  |  |  |  | 0.1 microF | 1.6k | 4Y | 999964 | 412 | 1591.65 | 20.86
2017-11-20 01:15 | 21.2 | 46.1 |  |  |  |  | 0.01 microF | 1.6k | 5Y | 1000024 | 498 | 1591.58 | 20.8
2017-11-20 01:20 | 21.2 | 46.1 |  |  |  |  | coax zero | 1.6k | 5Y | -1 | 2 | 1591.52 | 20
2017-11-20 01:25 | 21.2 | 46.1 |  |  |  |  | coax zero | 1.6k | 6Y | 0 | 14 | 1591.77 | 21.2
2017-11-20 01:30 | 21.2 | 45.9 |  |  |  |  | 1000 pF | 1.6k | 6Y | 999731 | -187 | 1591.75 | 21.18
2017-11-20 01:35 | 21.3 | 45.8 |  |  |  |  | GR1000A | 1.6k | 6Y | 999890 | -150 | 1591.67 | 20
2017-11-20 01:40 | 21.3 | 45.8 |  |  |  |  | GR1000B | 1.6k | 6Y | 999876 | -146 | 1591.6 | 20
2017-11-20 01:45 | 21.2 | 45.8 |  |  |  |  | AH11C | 1.6k | 6Y | 99985 | 104 | 1591.5 | 20
2017-11-20 01:50 | 21.2 | 45.8 |  |  |  |  | 100 pF | 1.6k | 6Y | 99990 | -15 | 1591.42 | 20.92
2017-11-20 01:55 | 21.2 | 45.8 |  |  |  |  | 100 pF | 1.6k | 7Y | 999946 | 1996 | 1591.32 | 20.91
2017-11-20 02:00 | 21.2 | 45.8 |  |  |  |  | AH11C | 1.6k | 7Y | 999901 | 2930 | 1591.19 | 20
2017-11-20 02:05 | 21.2 | 45.7 |  |  |  |  | coax zero | 1.6k | 7Y | 6 | 19 | 1591.12 | 20
2017-11-20 02:10 | 21.1 | 45.8 |  |  |  |  | AH11A | 1.6k | 7Y | 99995 | 201 | 1591.82 | 20
2017-11-20 02:15 | 21.2 | 45.8 |  |  |  |  | 10 pF | 1.6k | 7Y | 99987 | 210 | 1591.77 | 20.8
2017-11-20 02:20 | 21.3 | 45.4 |  |  |  |  | 1 pF | 1.6k | 7Y | 10004 | 46 | 1591.72 | 20.78
2017-11-20 02:25 | 21.3 | 45.4
2017-11-20 02:30 | 21.3 | 45.4
2017-11-20 02:35 | 21.3 | 45.3
2017-11-20 02:40 | 21.2 | 45.3
2017-11-20 02:45 | 21.2 | 45.3
2017-11-20 02:50 | 21.2 | 45.4
2017-11-20 02:55 | 21.2 | 45.3
2017-11-20 03:00 | 21.2 | 45.2
2017-11-20 03:05 | 21.2 | 45.2
2017-11-20 03:10 | 21.2 | 45.1
2017-11-20 03:15 | 21.3 | 45
2017-11-20 03:20 | 21.2 | 45
2017-11-20 03:25 | 21.2 | 45
2017-11-20 03:30 | 21.2 | 45.1
2017-11-20 03:35 | 21.2 | 44.9
2017-11-20 03:40 | 21.2 | 44.9
2017-11-20 03:45 | 21.3 | 44.8
2017-11-20 03:50 | 21.3 | 44.8
2017-11-20 03:55 | 21.2 | 44.9
2017-11-20 04:00 | 21.2 | 44.9
2017-11-20 04:05 | 21.3 | 44.7
2017-11-20 04:10 | 21.2 | 44.7
2017-11-20 04:15 | 21.2 | 44.6
2017-11-20 04:20 | 21.2 | 44.7
2017-11-20 04:25 | 21.2 | 44.6
2017-11-20 04:30 | 21.2 | 44.7
2017-11-20 04:35 | 21.2 | 44.7
2017-11-20 04:40 | 21.2 | 44.6
2017-11-20 04:45 | 21.2 | 44.5
2017-11-20 04:50 | 21.2 | 44.6
2017-11-20 04:55 | 21.2 | 44.5
2017-11-20 05:00 | 21.2 | 44.5
2017-11-20 05:05 | 21.2 | 44.4
2017-11-20 05:10 | 21.2 | 44.4
2017-11-20 05:15 | 21.2 | 44.4
2017-11-20 05:20 | 21.2 | 44.3
2017-11-20 05:25 | 21.2 | 44.4
2017-11-20 05:30 | 21.2 | 44.3
2017-11-20 05:35 | 21.3 | 44.2
2017-11-20 05:40 | 21.2 | 44.2
2017-11-20 05:45 | 21.3 | 44.2
2017-11-20 05:50 | 21.3 | 44.1
2017-11-20 05:55 | 21.3 | 44
2017-11-20 06:00 | 21.3 | 44
2017-11-20 06:05 | 21.3 | 43.9
2017-11-20 06:10 | 21.3 | 43.9
2017-11-20 06:15 | 21.2 | 44
2017-11-20 06:20 | 21.2 | 44
2017-11-20 06:25 | 21.3 | 43.8
2017-11-20 06:30 | 21.2 | 43.9
2017-11-20 06:35 | 21.2 | 43.9
2017-11-20 06:40 | 21.3 | 43.8
2017-11-20 06:45 | 21.2 | 43.8
2017-11-20 06:50 | 21.2 | 43.8
2017-11-20 06:55 | 21.3 | 43.7
2017-11-20 07:00 | 21.3 | 43.6
2017-11-20 07:05 | 21.3 | 43.6
2017-11-20 07:10 | 21.3 | 43.5
2017-11-20 07:15 | 21.3 | 43.5
2017-11-20 07:20 | 21.3 | 43.4
2017-11-20 07:25 | 21.3 | 43.5
2017-11-20 07:30 | 21.2 | 43.5
2017-11-20 07:35 | 21.3 | 43.4
2017-11-20 07:40 | 21.2 | 43.5
2017-11-20 07:45 | 21.3 | 43.3
2017-11-20 07:50 | 21.2 | 43.4
2017-11-20 07:55 | 21.3 | 43.3
2017-11-20 08:00 | 21.3 | 43.4
2017-11-20 08:05 | 21.3 | 43.4
2017-11-20 08:10 | 21.2 | 43.5
2017-11-20 08:15 | 21.3 | 43.4
2017-11-20 08:20 | 21.3 | 43.3
2017-11-20 08:25 | 21.2 | 43.5
2017-11-20 08:30 | 21.2 | 43.4
2017-11-20 08:35 | 21.2 | 43.5
2017-11-20 08:40 | 21.2 | 43.5
2017-11-20 08:45 | 21.3 | 43.5
2017-11-20 08:50 | 21.2 | 43.5
2017-11-20 08:55 | 21.2 | 43.5
2017-11-20 09:00 | 21.2 | 43.6
2017-11-20 09:05 | 21.3 | 43.6
2017-11-20 09:10 | 21.3 | 43.5
2017-11-20 09:15 | 21.2 | 43.7
2017-11-20 09:20 | 21.2 | 43.7
2017-11-20 09:25 | 21.2 | 43.8
2017-11-20 09:30 | 21.2 | 43.8
2017-11-20 09:35 | 21.2 | 43.8
2017-11-20 09:40 | 21.2 | 43.9
2017-11-20 09:45 | 21.2 | 44
2017-11-20 09:50 | 21.2 | 44.1
2017-11-20 09:55 | 21.2 | 44.2
2017-11-20 10:00 | 21.2 | 44.2
2017-11-20 10:05 | 21.3 | 44.2
2017-11-20 10:10 | 21.3 | 44.3
2017-11-20 10:15 | 21.3 | 44.4
2017-11-20 10:20 | 21.2 | 44.6
2017-11-20 10:25 | 21.2 | 44.8
2017-11-20 10:30 | 21.2 | 44.8
2017-11-20 10:35 | 21.2 | 45
2017-11-20 10:40 | 21.2 | 45
2017-11-20 10:45 | 21.3 | 45.1
2017-11-20 10:50 | 21.2 | 45.3
2017-11-20 10:55 | 21.2 | 45.4
2017-11-20 11:00 | 21.2 | 45.6
2017-11-20 11:05 | 21.2 | 45.7
2017-11-20 11:10 | 21.2 | 45.8
2017-11-20 11:15 | 21.2 | 46.1
2017-11-20 11:20 | 21.2 | 46.1
2017-11-20 11:25 | 21.1 | 46.5
2017-11-20 11:30 | 21.1 | 46.6
2017-11-20 11:35 | 21.1 | 46.5
2017-11-20 11:40 | 21.2 | 46.6
2017-11-20 11:45 | 21.1 | 46.9
2017-11-20 11:50 | 21.1 | 47
2017-11-20 11:55 | 21.1 | 47.3
2017-11-20 12:00 | 21.1 | 47.2
2017-11-20 12:05 | 21.1 | 47.4
2017-11-20 12:10 | 21.1 | 47.6
2017-11-20 12:15 | 21.1 | 47.7
2017-11-20 12:20 | 21.2 | 47.6
2017-11-20 12:25 | 21.2 | 47.7
2017-11-20 12:30 | 21.1 | 48.1
2017-11-20 12:35 | 21.1 | 47.9
2017-11-20 12:40 | 21.1 | 47.9
2017-11-20 12:45 | 21.1 | 48
2017-11-20 12:50 | 21.1 | 47.9
2017-11-20 12:55 | 21.1 | 48
2017-11-20 13:00 | 21.1 | 48.2
2017-11-20 13:05 | 21.1 | 48.2
2017-11-20 13:10 | 21.1 | 48
2017-11-20 13:15 | 21.2 | 48
2017-11-20 13:20 | 21.2 | 48
2017-11-20 13:25 | 21.2 | 48
2017-11-20 13:30 | 21.2 | 48
2017-11-20 13:35 | 21.1 | 48.1
2017-11-20 13:40 | 21.2 | 48.2
2017-11-20 13:45 | 21.1 | 48.3
2017-11-20 13:50 | 21.1 | 48.4
2017-11-20 13:55 | 21 | 48.6
2017-11-20 14:00 | 21.1 | 48.6
2017-11-20 14:05 | 21.1 | 48.6
2017-11-20 14:10 | 21.1 | 48.7
2017-11-20 14:15 | 21.1 | 48.7
2017-11-20 14:20 | 21.1 | 48.8
2017-11-20 14:25 | 21.1 | 48.9
2017-11-20 14:30 | 21 | 48.9
2017-11-20 14:35 | 21.1 | 48.9
2017-11-20 14:40 | 21.1 | 49
2017-11-20 14:45 | 21.2 | 49
2017-11-20 14:50 | 21.1 | 49.6
2017-11-20 14:55 | 21.2 | 49.3
2017-11-20 15:00 | 21.1 | 49.4
2017-11-20 15:05 | 21.1 | 49.4
2017-11-20 15:10 | 21.1 | 49.6
2017-11-20 15:15 | 21.2 | 49.3
2017-11-20 15:20 | 21.2 | 49
2017-11-20 15:25 | 21.2 | 48.9
2017-11-20 15:30 | 21.2 | 48.8
2017-11-20 15:35 | 21.3 | 48.7
2017-11-20 15:40 | 21.2 | 48.6
2017-11-20 15:45 | 21.3 | 48.5
2017-11-20 15:50 | 21.2 | 48.6
2017-11-20 15:55 | 21.2 | 48.7
2017-11-20 16:00 | 21.2 | 48.8
2017-11-20 16:05 | 21.2 | 48.9
2017-11-20 16:10 | 21.1 | 48.9
2017-11-20 16:15 | 21.1 | 48.8
2017-11-20 16:20 | 21.1 | 48.7
2017-11-20 16:25 | 21.1 | 48.7
2017-11-20 16:30 | 21.2 | 48.6
2017-11-20 16:35 | 21.1 | 48.7
2017-11-20 16:40 | 21.1 | 48.6
2017-11-20 16:45 | 21.1 | 48.6
2017-11-20 16:50 | 21.1 | 48.6
2017-11-20 16:55 | 21.1 | 48.6
2017-11-20 17:00 | 21.1 | 48.5
2017-11-20 17:05 | 21 | 48.7
2017-11-20 17:10 | 21.1 | 48.7
2017-11-20 17:15 | 21.1 | 48.7
2017-11-20 17:20 | 21 | 48.8
2017-11-20 17:25 | 21 | 48.9
2017-11-20 17:30 | 21 | 48.9
2017-11-20 17:35 | 21 | 48.9
2017-11-20 17:40 | 21.1 | 48.7
2017-11-20 17:45 | 21 | 48.8
2017-11-20 17:50 | 21 | 48.8
2017-11-20 17:55 | 21 | 48.8
2017-11-20 18:00 | 21 | 48.8
2017-11-20 18:05 | 21 | 48.7
2017-11-20 18:10 | 21 | 48.6
2017-11-20 18:15 | 21 | 48.5
2017-11-20 18:20 | 21 | 48.6
2017-11-20 18:25 | 21 | 48.5
2017-11-20 18:30 | 21.1 | 48.4
2017-11-20 18:35 | 21 | 48.4
2017-11-20 18:40 | 21 | 48.4
2017-11-20 18:45 | 21 | 48.5
2017-11-20 18:50 | 21.1 | 48.4
2017-11-20 18:55 | 21 | 48.5
2017-11-20 19:00 | 21 | 48.5
2017-11-20 19:05 | 21.1 | 48.3
2017-11-20 19:10 | 21 | 48.3
2017-11-20 19:15 | 21 | 48.4
2017-11-20 19:20 | 21 | 48.5
2017-11-20 19:25 | 21 | 48.5
2017-11-20 19:30 | 21 | 48.5
2017-11-20 19:35 | 21 | 48.4
2017-11-20 19:40 | 21 | 48.4
2017-11-20 19:45 | 21 | 48.5
2017-11-20 19:50 | 21 | 48.4
2017-11-20 19:55 | 21 | 48.5
2017-11-20 20:00 | 21 | 48.4
2017-11-20 20:05 | 21 | 48.4
2017-11-20 20:10 | 21 | 48.5
2017-11-20 20:15 | 21 | 48.5
2017-11-20 20:20 | 21 | 48.5
2017-11-20 20:25 | 21 | 48.5
2017-11-20 20:30 | 21 | 48.5
2017-11-20 20:35 | 21 | 48.6
2017-11-20 20:40 | 20.9 | 48.6
2017-11-20 20:45 | 20.9 | 48.6
2017-11-20 20:50 | 21 | 48.5
2017-11-20 20:55 | 20.9 | 48.6
2017-11-20 21:00 | 21 | 48.4
2017-11-20 21:05 | 21 | 48.4
2017-11-20 21:10 | 21 | 48.3
2017-11-20 21:15 | 21 | 48.3
2017-11-20 21:20 | 21.1 | 48.2
2017-11-20 21:25 | 21 | 48.3
2017-11-20 21:30 | 21 | 48.3
2017-11-20 21:35 | 21.1 | 48.2
2017-11-20 21:40 | 21 | 48.2
2017-11-20 21:45 | 21 | 48.2
2017-11-20 21:50 | 21 | 48.2
2017-11-20 21:55 | 21 | 48.3
2017-11-20 22:00 | 20.9 | 48.3
2017-11-20 22:05 | 21 | 48.2
2017-11-20 22:10 | 20.9 | 48.4
2017-11-20 22:15 | 21 | 48.1
2017-11-20 22:20 | 21 | 48.2
2017-11-20 22:25 | 21 | 48.1
2017-11-20 22:30 | 21 | 48.1
2017-11-20 22:35 | 21 | 48.1
2017-11-20 22:40 | 21 | 48
2017-11-20 22:45 | 21 | 48
2017-11-20 22:50 | 21 | 48
2017-11-20 22:55 | 21 | 47.9
2017-11-20 23:00 | 21 | 47.8
2017-11-20 23:05 | 21 | 47.7
2017-11-20 23:10 | 21 | 47.6
2017-11-20 23:15 | 21.1 | 47.6
2017-11-20 23:20 | 21 | 47.6
2017-11-20 23:25 | 21 | 47.6
2017-11-20 23:30 | 21 | 47.7
2017-11-20 23:35 | 21 | 47.5
2017-11-20 23:40 | 21 | 47.5
2017-11-20 23:45 | 21.1 | 47.4
2017-11-20 23:50 | 21 | 47.4
2017-11-20 23:55 | 21 | 47.5
2017-11-21 | 21 | 47.3
2017-11-21 00:05 | 21 | 47.3
2017-11-21 00:10 | 21 | 47.2
2017-11-21 00:15 | 21 | 47.2
2017-11-21 00:20 | 21 | 47.1
2017-11-21 00:25 | 21 | 47.1
2017-11-21 00:30 | 21 | 47
2017-11-21 00:35 | 21 | 47
2017-11-21 00:40 | 21 | 47
2017-11-21 00:45 | 21 | 47
2017-11-21 00:50 | 21 | 46.8
2017-11-21 00:55 | 21.1 | 46.7
2017-11-21 01:00 | 21 | 46.7
2017-11-21 01:05 | 21 | 46.7
2017-11-21 01:10 | 21 | 46.7
2017-11-21 01:15 | 21 | 46.7
2017-11-21 01:20 | 21 | 46.6
2017-11-21 01:25 | 21 | 46.6
2017-11-21 01:30 | 21 | 46.6
2017-11-21 01:35 | 21 | 46.5
2017-11-21 01:40 | 21 | 46.5
2017-11-21 01:45 | 21 | 46.5
2017-11-21 01:50 | 21 | 46.5
2017-11-21 01:55 | 21 | 46.4
2017-11-21 02:00 | 21 | 46.3
2017-11-21 02:05 | 21 | 46.3
2017-11-21 02:10 | 21 | 46.2
2017-11-21 02:15 | 21 | 46.3
2017-11-21 02:20 | 21 | 46.3
2017-11-21 02:25 | 21 | 46.2
2017-11-21 02:30 | 21.1 | 46.1
2017-11-21 02:35 | 21.1 | 46.1
2017-11-21 02:40 | 21.1 | 46.1
2017-11-21 02:45 | 21 | 46.1
2017-11-21 02:50 | 21 | 46.2
2017-11-21 02:55 | 21 | 46.1
2017-11-21 03:00 | 21.1 | 46
2017-11-21 03:05 | 21 | 46.1
2017-11-21 03:10 | 21 | 46.1
2017-11-21 03:15 | 21.1 | 46.1
2017-11-21 03:20 | 21 | 46.1
2017-11-21 03:25 | 21.1 | 46.1
2017-11-21 03:30 | 21.1 | 46.1
2017-11-21 03:35 | 21 | 46.1
2017-11-21 03:40 | 21 | 46.2
2017-11-21 03:45 | 21 | 46.3
2017-11-21 03:50 | 21 | 46.3
2017-11-21 03:55 | 21 | 46.4
2017-11-21 04:00 | 21.1 | 46.2
2017-11-21 04:05 | 21 | 46.3
2017-11-21 04:10 | 21 | 46.4
2017-11-21 04:15 | 21 | 46.4
2017-11-21 04:20 | 21 | 46.3
2017-11-21 04:25 | 21 | 46.3
2017-11-21 04:30 | 21 | 46.2
2017-11-21 04:35 | 21 | 46.3
2017-11-21 04:40 | 21 | 46.2
2017-11-21 04:45 | 21 | 46.3
2017-11-21 04:50 | 21 | 46.3
2017-11-21 04:55 | 21 | 46.3
2017-11-21 05:00 | 21 | 46.3
2017-11-21 05:05 | 21 | 46.2
2017-11-21 05:10 | 21 | 46.2
2017-11-21 05:15 | 21 | 46.2
2017-11-21 05:20 | 21 | 46.1
2017-11-21 05:25 | 21 | 46.1
2017-11-21 05:30 | 21 | 46.2
2017-11-21 05:35 | 21 | 46.1
2017-11-21 05:40 | 21 | 46.2
2017-11-21 05:45 | 21 | 46.1
2017-11-21 05:50 | 21 | 46
2017-11-21 05:55 | 21 | 46.1
2017-11-21 06:00 | 21 | 46
2017-11-21 06:05 | 21 | 46.1
2017-11-21 06:10 | 21 | 46
2017-11-21 06:15 | 21 | 45.9
2017-11-21 06:20 | 21 | 45.9
2017-11-21 06:25 | 21 | 45.8
2017-11-21 06:30 | 21.1 | 45.7
2017-11-21 06:35 | 21 | 45.9
2017-11-21 06:40 | 21 | 45.7
2017-11-21 06:45 | 21 | 45.8
2017-11-21 06:50 | 21 | 45.8
2017-11-21 06:55 | 21.1 | 45.6
2017-11-21 07:00 | 21 | 45.7
2017-11-21 07:05 | 21 | 45.7
2017-11-21 07:10 | 21 | 45.8
2017-11-21 07:15 | 21.1 | 45.6
2017-11-21 07:20 | 21 | 45.6
2017-11-21 07:25 | 21 | 45.6
2017-11-21 07:30 | 21.1 | 45.6
2017-11-21 07:35 | 21 | 45.7
2017-11-21 07:40 | 21.1 | 45.5
2017-11-21 07:45 | 21 | 45.6
2017-11-21 07:50 | 21 | 45.6
2017-11-21 07:55 | 21 | 45.7
2017-11-21 08:00 | 21 | 45.6
2017-11-21 08:05 | 21 | 45.6
2017-11-21 08:10 | 21 | 45.5
2017-11-21 08:15 | 21.1 | 45.5
2017-11-21 08:20 | 21 | 45.6
2017-11-21 08:25 | 21.1 | 45.6
2017-11-21 08:30 | 21.1 | 45.5
2017-11-21 08:35 | 21.1 | 45.6
2017-11-21 08:40 | 21 | 45.7
2017-11-21 08:45 | 21.1 | 45.6
2017-11-21 08:50 | 21.1 | 45.7
2017-11-21 08:55 | 21.1 | 45.6
2017-11-21 09:00 | 21 | 45.9
2017-11-21 09:05 | 21 | 45.9
2017-11-21 09:10 | 21 | 45.9
2017-11-21 09:15 | 21 | 46
2017-11-21 09:20 | 21 | 46
2017-11-21 09:25 | 21.1 | 46
2017-11-21 09:30 | 21 | 46.2
2017-11-21 09:35 | 21.1 | 46.1
2017-11-21 09:40 | 21.1 | 46.1
2017-11-21 09:45 | 21 | 46.2
2017-11-21 09:50 | 21.1 | 46.1
2017-11-21 09:55 | 21 | 46.3
2017-11-21 10:00 | 21 | 46.4
2017-11-21 10:05 | 21 | 46.5
2017-11-21 10:10 | 21 | 46.4
2017-11-21 10:15 | 21 | 46.5
2017-11-21 10:20 | 21 | 46.4
2017-11-21 10:25 | 21 | 46.5
2017-11-21 10:30 | 21.1 | 46.4
2017-11-21 10:35 | 21 | 46.6
2017-11-21 10:40 | 21.1 | 46.6
2017-11-21 10:45 | 21.1 | 46.6
2017-11-21 10:50 | 21.1 | 46.7
2017-11-21 10:55 | 21 | 46.9
2017-11-21 11:00 | 21.1 | 46.8
2017-11-21 11:05 | 21.1 | 46.9
2017-11-21 11:10 | 21 | 47
2017-11-21 11:15 | 21 | 47.1
2017-11-21 11:20 | 21 | 47.1
2017-11-21 11:25 | 21.1 | 47.1
2017-11-21 11:30 | 21 | 47.3
2017-11-21 11:35 | 21 | 47.3
2017-11-21 11:40 | 21 | 47.4
2017-11-21 11:45 | 21 | 47.5
2017-11-21 11:50 | 21 | 47.3
2017-11-21 11:55 | 21.1 | 47.4
2017-11-21 12:00 | 21.1 | 47.4
2017-11-21 12:05 | 21 | 47.5
2017-11-21 12:10 | 21 | 47.6
2017-11-21 12:15 | 21 | 47.6
2017-11-21 12:20 | 21 | 47.7
2017-11-21 12:25 | 21 | 47.6
2017-11-21 12:30 | 21 | 47.6
2017-11-21 12:35 | 21.1 | 47.6
2017-11-21 12:40 | 21.1 | 47.7
2017-11-21 12:45 | 21.1 | 47.6
2017-11-21 12:50 | 21.1 | 47.7
2017-11-21 12:55 | 21.1 | 47.7
2017-11-21 13:00 | 21.1 | 47.8
2017-11-21 13:05 | 21.1 | 47.8
2017-11-21 13:10 | 21 | 47.8
2017-11-21 13:15 | 21 | 47.9
2017-11-21 13:20 | 21 | 47.9
2017-11-21 13:25 | 21 | 48
2017-11-21 13:30 | 21 | 48
2017-11-21 13:35 | 21 | 48.2
2017-11-21 13:40 | 21 | 48.1
2017-11-21 13:45 | 21 | 48.2
2017-11-21 13:50 | 21 | 48.3
2017-11-21 13:55 | 21 | 48.5
2017-11-21 14:00 | 21 | 48.7
2017-11-21 14:05 | 20.9 | 49
2017-11-21 14:10 | 20.9 | 49.2
2017-11-21 14:15 | 20.9 | 49.4
2017-11-21 14:20 | 21 | 49.2
2017-11-21 14:25 | 21 | 49.3
2017-11-21 14:30 | 20.9 | 49.5
2017-11-21 14:35 | 20.9 | 49.8
2017-11-21 14:40 | 20.9 | 49.8
2017-11-21 14:45 | 20.9 | 49.8
2017-11-21 14:50 | 20.9 | 49.8
2017-11-21 14:55 | 21 | 49.7
2017-11-21 15:00 | 20.9 | 49.7
2017-11-21 15:05 | 20.9 | 49.9
2017-11-21 15:10 | 21 | 49.8
2017-11-21 15:15 | 20.9 | 49.9
2017-11-21 15:20 | 21 | 50
2017-11-21 15:25 | 21 | 49.9
2017-11-21 15:30 | 21 | 49.9
2017-11-21 15:35 | 20.9 | 50.1
2017-11-21 15:40 | 20.9 | 50.1
2017-11-21 15:45 | 20.9 | 50.1
2017-11-21 15:50 | 20.9 | 50.3
2017-11-21 15:55 | 20.9 | 50.2
2017-11-21 16:00 | 20.9 | 50.2
2017-11-21 16:05 | 20.9 | 50.2
2017-11-21 16:10 | 20.9 | 50.3
2017-11-21 16:15 | 20.9 | 50.2
2017-11-21 16:20 | 20.9 | 50.4
2017-11-21 16:25 | 20.9 | 50.4
2017-11-21 16:30 | 20.9 | 50.4
2017-11-21 16:35 | 20.9 | 50.4
2017-11-21 16:40 | 20.9 | 50.4
2017-11-21 16:45 | 20.9 | 50.5
2017-11-21 16:50 | 20.9 | 50.4
2017-11-21 16:55 | 20.9 | 50.4
2017-11-21 17:00 | 20.9 | 50.4
2017-11-21 17:05 | 20.9 | 50.3
2017-11-21 17:10 | 21 | 50.2
2017-11-21 17:15 | 20.9 | 50.3
2017-11-21 17:20 | 20.9 | 50.3
2017-11-21 17:25 | 20.9 | 50.2
2017-11-21 17:30 | 20.9 | 50.2
2017-11-21 17:35 | 20.9 | 50.2
2017-11-21 17:40 | 20.9 | 50.3
2017-11-21 17:45 | 20.9 | 50.1
2017-11-21 17:50 | 20.9 | 50.1
2017-11-21 17:55 | 20.9 | 50.1
2017-11-21 18:00 | 20.9 | 50.3
2017-11-21 18:05 | 21 | 50.2
2017-11-21 18:10 | 20.9 | 50.3
2017-11-21 18:15 | 20.9 | 50.4
2017-11-21 18:20 | 20.9 | 50.4
2017-11-21 18:25 | 21 | 50.3
2017-11-21 18:30 | 20.9 | 50.3
2017-11-21 18:35 | 21 | 50.2
2017-11-21 18:40 | 21 | 50.1
2017-11-21 18:45 | 20.9 | 50.3
2017-11-21 18:50 | 21 | 50.2
2017-11-21 18:55 | 20.9 | 50.2
2017-11-21 19:00 | 20.9 | 50.2
2017-11-21 19:05 | 21 | 50
2017-11-21 19:10 | 21 | 50
2017-11-21 19:15 | 21 | 50
2017-11-21 19:20 | 20.9 | 50.1
2017-11-21 19:25 | 21 | 50
2017-11-21 19:30 | 21 | 50.1
2017-11-21 19:35 | 21 | 50
2017-11-21 19:40 | 21 | 49.8
2017-11-21 19:45 | 21 | 49.7
2017-11-21 19:50 | 21 | 49.7
2017-11-21 19:55 | 21 | 49.6
2017-11-21 20:00 | 21 | 49.5
2017-11-21 20:05 | 21 | 49.6
2017-11-21 20:10 | 21 | 49.4
2017-11-21 20:15 | 21 | 49.4
2017-11-21 20:20 | 21 | 49.4
2017-11-21 20:25 | 21 | 49.4
2017-11-21 20:30 | 20.9 | 49.4
2017-11-21 20:35 | 21 | 49.2
2017-11-21 20:40 | 21 | 49.3
2017-11-21 20:45 | 20.9 | 49.4
2017-11-21 20:50 | 21 | 49.3
2017-11-21 20:55 | 21 | 49.3
2017-11-21 21:00 | 21 | 49.3
2017-11-21 21:05 | 21 | 49.3
2017-11-21 21:10 | 21 | 49.4
2017-11-21 21:15 | 21 | 49.4
2017-11-21 21:20 | 21 | 49.4
2017-11-21 21:25 | 20.9 | 49.5
2017-11-21 21:30 | 20.9 | 49.5
2017-11-21 21:35 | 21 | 49.4
2017-11-21 21:40 | 21 | 49.4
2017-11-21 21:45 | 21 | 49.4
2017-11-21 21:50 | 21 | 49.4
2017-11-21 21:55 | 20.9 | 49.5
2017-11-21 22:00 | 21 | 49.4
2017-11-21 22:05 | 21 | 49.3
2017-11-21 22:10 | 21 | 49.4
2017-11-21 22:15 | 21 | 49.4
2017-11-21 22:20 | 21 | 49.4
2017-11-21 22:25 | 21 | 49.3
2017-11-21 22:30 | 21 | 49.4
2017-11-21 22:35 | 21 | 49.4
2017-11-21 22:40 | 21 | 49.3
2017-11-21 22:45 | 21 | 49.4
2017-11-21 22:50 | 21 | 49.4
2017-11-21 22:55 | 21 | 49.4
2017-11-21 23:00 | 21 | 49.3
2017-11-21 23:05 | 21 | 49.4
2017-11-21 23:10 | 20.9 | 49.5
2017-11-21 23:15 | 20.9 | 49.5
2017-11-21 23:20 | 20.9 | 49.5
2017-11-21 23:25 | 20.9 | 49.5
2017-11-21 23:30 | 21 | 49.4
2017-11-21 23:35 | 21 | 49.4
2017-11-21 23:40 | 21 | 49.4
2017-11-21 23:45 | 21 | 49.4
2017-11-21 23:50 | 21 | 49.3
2017-11-21 23:55 | 21 | 49.2

## Results
Template for decade and individual capacitors measured on Universal Bridge | [8 empty] | Checked by:
Capacitors | Babcock HP16380 |  |  | Client | Babcock
Date | 2017-11-20 |  |  | Job number | 96330108 | [6 empty] | 20.77
|  |  |  | S Number | S22013 | [6 empty] | 0.091378
[9 empty] | Nominal | loss angle |  | GMH5I readout
Item | Label | Range | a | b | Capacitance |  | Conductance |  | Frequency | milliradians | differentzero | temperature
Hz
1 microF | 1 microF3Y | 3Y | 999846 | 1716 | 0.99992 | microF | 0.0017179 | milliS | 1591.5 | 0.17181
coax zero | coax zero3Y | 3Y | 0 | -2 | 0 | microF | 0 | milliS | 1591.5 | #DIV/0!
coax zero | coax zero4Y | 4Y | 0 | -1 | 0 | nF | 0 | microS | 1591.5 | #DIV/0!
0.1 microF | 0.1 microF4Y | 4Y | 999964 | 412 | 100.0039 | nF | 0.041298 | microS | 1591.5 | 0.041297 |  | 20.68
0.01 microF | 0.01 microF5Y | 5Y | 1000024 | 498 | 10.001 | nF | 0.0049597 | microS | 1591.5 | 0.049594
coax zero | coax zero5Y | 5Y | -1 | 2 | 0 | nF | 0 | microS | 1591.5 | #DIV/0! |  | 20.81
coax zero | coax zero6Y | 6Y | 0 | 14 | 0 | pF | 0 | nS | 1591.5 | #DIV/0!
1000 pF | 1000 pF6Y | 6Y | 999731 | -187 | 999.8688 | pF | -0.201 | nS | 1591.5 | -0.020103 |  | 20.73
GR1000A | GR1000A6Y | 6Y | 999890 | -150 | 1000.0279 | pF | -0.164 | nS | 1591.5 | -0.0164
GR1000B | GR1000B6Y | 6Y | 999876 | -146 | 1000.0138 | pF | -0.16 | nS | 1591.5 | -0.016 |  | 20.72
AH11C | AH11C6Y | 6Y | 99985 | 104 | 99.9988 | pF | 0.09 | nS | 1591.5 | 0.090004
100 pF | 100 pF6Y | 6Y | 99990 | -15 | 100.0038 | pF | -0.029 | nS | 1591.5 | -0.029 |  | 20.91
100 pF | 100 pF7Y | 7Y | 999946 | 1996 | 100.0003 | pF | 0.19769 | nS | 1591.5 | 0.19769
AH11C | AH11C7Y | 7Y | 999901 | 2930 | 99.9958 | pF | 0.29108 | nS | 1591.5 | 0.2911
coax zero | coax zero7Y | 7Y | 6 | 19 | 0 | pF | 0 | nS | 1591.5 | #DIV/0! |  | 20.77
AH11A | AH11A7Y | 7Y | 99995 | 201 | 9.9995 | pF | 0.018199 | nS | 1591.5 | 0.182 |  | 0.091378
10 pF | 10 pF7Y | 7Y | 99987 | 210 | 9.9987 | pF | 0.019099 | nS | 1591.5 | 0.19102
1 pF | 1 pF7Y | 7Y | 10004 | 46 | 0.99986 | pF | 0.0026998 | nS | 1591.5 | 0.27003
[5 empty] | #N/A | #N/A | #N/A | #N/A | 1591.5 | #N/A
[5 empty] | #N/A | #N/A | #N/A | #N/A | 1591.5 | #N/A
[5 empty] | #N/A | #N/A | #N/A | #N/A | 1591.5 | #N/A
[5 empty] | #N/A | #N/A | #N/A | #N/A | 1591.5 | #N/A
[5 empty] | #N/A | #N/A | #N/A | #N/A | 1591.5 | #N/A
[5 empty] | #N/A | #N/A | #N/A | #N/A | 1591.5 | #N/A
[5 empty] | #N/A | #N/A | #N/A | #N/A | 1591.5 | #N/A
[5 empty] | #N/A | #N/A | #N/A | #N/A | 1591.5 | #N/A
[5 empty] | #N/A | #N/A | #N/A | #N/A | 1591.5 | #N/A
[5 empty] | #N/A | #N/A | #N/A | #N/A | 1591.5 | #N/A
[5 empty] | #N/A | #N/A | #N/A | #N/A | 1591.5 | #N/A
[5 empty] | #N/A | #N/A | #N/A | #N/A | 1591.5 | #N/A
[5 empty] | #N/A | #N/A | #N/A | #N/A | 1591.5 | #N/A
[5 empty] | #N/A | #N/A | #N/A | #N/A | 1591.5 | #N/A
[5 empty] | #N/A | #N/A | #N/A | #N/A | 1591.5 | #N/A
[5 empty] | #N/A | #N/A | #N/A | #N/A | 1591.5 | #N/A
[5 empty] | #N/A | #N/A | #N/A | #N/A | 1591.5 | #N/A
[5 empty] | #N/A | #N/A | #N/A | #N/A | 1591.5 | #N/A
[5 empty] | #N/A | #N/A | #N/A | #N/A | 1591.5 | #N/A
[5 empty] | #N/A | #N/A | #N/A | #N/A | 1591.5 | #N/A
[5 empty] | #N/A | #N/A | #N/A | #N/A | 1591.5 | #N/A
[5 empty] | #N/A | #N/A | #N/A | #N/A | 1591.5 | #N/A
[5 empty] | #N/A | #N/A | #N/A | #N/A | 1591.5 | #N/A
[5 empty] | #N/A | #N/A | #N/A | #N/A | 1591.5 | #N/A
[5 empty] | #N/A | #N/A | #N/A | #N/A | 1591.5 | #N/A
[5 empty] | #N/A | #N/A | #N/A | #N/A | 1591.5 | #N/A
[5 empty] | #N/A | #N/A | #N/A | #N/A | 1591.5 | #N/A
[5 empty] | #N/A | #N/A | #N/A | #N/A | 1591.5 | #N/A
[5 empty] | #N/A | #N/A | #N/A | #N/A | 1591.6 | #N/A
[5 empty] | #N/A | #N/A | #N/A | #N/A | 1591.4 | #N/A
[5 empty] | #N/A | #N/A | #N/A | #N/A | 1591.4 | #N/A
[5 empty] | #N/A | #N/A | #N/A | #N/A | 1591.4 | #N/A
[5 empty] | #N/A | #N/A | #N/A | #N/A | 1591.7 | #N/A
[5 empty] | #N/A | #N/A | #N/A | #N/A | 1591.7 | #N/A
[5 empty] | #N/A | #N/A | #N/A | #N/A | 1591.7 | #N/A
[5 empty] | #N/A | #N/A | #N/A | #N/A | 1591.7 | #N/A
[5 empty] | #N/A | #N/A | #N/A | #N/A | 1591.7 | #N/A
[5 empty] | #N/A | #N/A | #N/A | #N/A | 1591.7 | #N/A
[5 empty] | #N/A | #N/A | #N/A | #N/A | 1591.7 | #N/A
[5 empty] | #N/A | #N/A | #N/A | #N/A | 1591.7 | #N/A
[5 empty] | #N/A | #N/A | #N/A | #N/A | 1591.7 | #N/A
[5 empty] | #N/A | #N/A | #N/A | #N/A | 1591.7 | #N/A
[5 empty] | #N/A | #N/A | #N/A | #N/A | 1591.6 | #N/A

## Universal Bridge
New UB calibrations need to be entered in Data block and accessed through the Date |  |  | This table is the usual form for calculating spreadsheet corrections | [6 empty] | Date | C1 | R4A | R4B | R4C | G1 | G2 | c | d
|  |  | Component | Value | Date |  |  |  | 1.6 kHz | 1985-02-14 | 32 | 30.1 | -12 | 53 | 13 | 8.7 | <10 | <10 | except for range 4, up to 44? |  |  |  | -91.32 | 0 | 0 | 0 | 0 | 0
|  |  |  | (ppm) | 2017-11-10 |  |  |  | DC for resistors | 1988-11-30 | 123.32 | 35 | -12 | #N/A | 9 | 4 | 0 | 0 |  |  |  |  | 0 | 4.9 | 0 | #N/A | -4 | -4.7 | 0
Date | Record | Comments | C_1 | 149 | UBIII |  |  |  | DC for resistors | 1991-03-12 | 133 | 37 | -14 | 61 | 10 | 6 | 0 | 0 |  |  |  |  | 9.68 | 6.9 | -2 | 8 | -3 | -2.7 | 0
1981-05-22 | PEL Meas V | Check Y&Z ranges 6 & 4 @ 400 Hz | R_4A | 54.3489 | UBIII |  |  |  | DC for resistors | 1995-08-22 | 127.406 | 40.34 | -17.21 | 52 | 11.21 | 6.21 | 0 | 0 |  |  |  |  | 4.086 | 10.24 | -5.21 | -1 | -1.79 | -2.49 | 0
1983-12-15 | Universal Bridge II | Full calibration of standards and of all dial ratios and performance checks. | R_4B | -8.2635 | UBIII |  |  |  | DC for resistors | 1998-11-03 | 135.27 | 43.3741 | -16.3523 | 67.0039 | 11.1844 | 6.6891 | 0 | 0 |  |  |  |  | 11.95 | 13.2741 | -4.3523 | 14.0039 | -1.8156 | -2.0109 | 0
1984-02-10 |  | completion of above | R_4C | 66.7468 | UBIII |  |  |  | DC for resistors | 2001-09-08 | 124.1565 | 46.006 | -17.3198 | 64.091 | 12.0228 | 7.4763 | 0 | 0 |  |  |  |  | 0.83653 | 15.906 | -5.3198 | 11.091 | -0.97725 | -1.2237 | 0
1988-11-30 | UB II | component values measured | G_1 | 13.822 | UBIII |  |  |  | DC for resistors | 2004-01-22 | 131.7148 | 47.52 | -16.236 | 63.74 | 12.616 | 7.596 | 0 | 0 |  |  |  |  | 8.3948 | 17.42 | -4.236 | 10.74 | -0.384 | -1.104 | 0
1991-03-12 | UB II | component values measured | G_2 | 8.4272 | UBIII |  |  |  | DC for resistors | 2004-02-18 | 123.9 | 47.52 | -16.236 | 63.74 | 12.616 | 7.596 | 5.2 | 0 |  |  |  |  | 0.58 | 17.42 | -4.236 | 10.74 | -0.384 | -1.104 | 5.2
1991-05-16 | UB II | problem with '6' dial, repaired! | [6 empty] | DC for resistors | 2005-02-18 | 129.4493 | 46.54 | -16.96 | 61.5 | 13.84 | 9 | 5.7 | 0 |  |  |  |  | 6.1293 | 16.44 | -4.96 | 8.5 | 0.84 | 0.3 | 5.7
1991-06-18 | AI III | problem with 0.01 mF on range 5Y, lack of DC path for current pump, AVOID | Z3=(R4/G2)(aG1+ßjwC1)(1+c+jd) |  |  | assume negligible c,d |  |  | DC for resistors | 2009-08-13 | 130.2987 | 50.8747 | -14.1043 | 64.5247 | 12.5073 | 7.4753 | 2 | 0 |  |  |  |  | 6.9787 | 20.7747 | -2.1043 | 11.5247 | -0.4927 | -1.2247 | 2
1995-08-22 | UB II | bridge calibration | where Z3=R+jwC | correction_c | 10.97 |  |  |  | DC for resistors | 2014-07-07 | 143.1 | 54.1261 | -9.8547 | 67.1933 | 12.9485 | 7.4961 | 10.97 | 0 |  |  |  |  | 19.78 | 24.0261 | 2.1453 | 14.1933 | -0.051497 | -1.2039 | 10.97
1998-11-03 | UB II | bridge calibration, using HP3458 as transfer instrument | Y4=(G3/G2)(aG1+ßjwC1)(1+c+jd) |  |  | (G3=1/R4*) |  |  | C1 jumped? | 2015-05-01 | 162.9971 | 54.1261 | -9.8547 | 67.1933 | 12.9485 | 7.4961 | 10.97 | 0 |  |  |  |  | 39.6771
2001-09-08 | UBII | bridge calibration, using HP3458 as transfer instrument: Range 5Y problem fixed | where Y4=G+jwC | [5 empty] | C1 settled | 2016-09-01 | 140.1 | 54.1261 | -9.8547 | 67.1933 | 12.9485 | 7.4961 | 10.97 | 0
2004-01-22 | UBII | bridge calibration, using HP3458 as transfer instrument: Range 5Y problem REALLY fixed | [6 empty] | partial calibration | 2017-11-10 | 149 | 54.3489 | -8.2635 | 66.7468 | 13.822 | 8.4272 | 10.97 | 0
2004-02-18 | UBII | remeasure capacitor, same dc resistance values as for 22/01/04 | R4A is used for ranges 5 to 1 inclusive | [5 empty] | UB1000 jump | AH2700 readings
2005-02-18 | UBII | bridge calibration, following MSLT.E.023.004. Uncertainties have increased. | R4B is used on range 6 only | [5 empty] | 2014-07-08 | 1000.1449
2009-08-13 | UBIII | bridge calibration, following MSLT.E.023.004/5. Uncertainties have increased slightly from 2005. | R4C is used on range 7 only | [5 empty] | 2015-05-01 | 1000.1648 | [9 empty] | AH2700A | ppm | mrad
2014-07-07 | UBIII | bridge calibration, following MSLT.E.023.006. Uncertainties unchanged from 2009. | [6 empty] | increase | 19.8971 | ppm | Zeros | Range | Differentzero | Range Label | a | b |  |  | R_4C | 15.0002 | 0.096
2015-05-01 | UBIII | shift in C1, likely due to shifting of UB to power lab while air con ducting being worked on. | [10 empty] | 3Y |  | 3Y | 0 | -2 | coaxial zero
2016-09-01 | KJLabDiary1 | C1 appears to have recovered | [10 empty] | 4Y |  | 4Y | 0 | -1 | coaxial zero
2017-11-10 | UBIII | Partial recalibration | [10 empty] | 5Y |  | 5Y | -1 | 2 | coaxial zero
[13 empty] | 6Y |  | 6Y | 0 | 14 | coaxial zero
|  |  |  | Range | a correction | b correction | [6 empty] | 7Y |  | 7Y | 6 | 19 | coaxial zero
[5 empty] | (ppm) | (ppm)
|  |  |  | 1Z | 205.8917 | 70.7137
|  |  |  | 2Z | 205.8917 | 70.7137
|  |  |  | 3Z | 205.8917 | 70.7137
|  |  |  | 4Z | 205.8917 | 70.7137
|  |  |  | 5Z | 205.8917 | 70.7137
|  |  |  | 6Z | 143.2793 | 8.1013
|  |  |  | 7Z | 218.2896 | 83.1116
|  |  |  | 1Y | 75.2539 | -59.9241
|  |  |  | 2Y | 75.2539 | -59.9241
|  |  |  | 3Y | 75.2539 | -59.9241
|  |  |  | 4Y | 75.2539 | -59.9241
|  |  |  | 5Y | 75.2539 | -59.9241
|  |  |  | 6Y | 137.8663 | 2.6882
|  |  |  | 7Y | 62.856 | -72.322
[19 empty] | Symbol font versions |  | Plain version
[5 empty] | Multiplier (C&L) | Unit Name | Unit Multiplier | Multiplier (G&R) | Unit Name | Unit Multiplier | CMC n for L | CMC m for R | [6 empty] | Unit Name | Unit Name | Unit Name | Unit Name
|  |  |  | 1Z | 0.0001 | microH | 1e-06 | 0.0001 | milliohm | 0.001 | 0.2 | 0.002 | [6 empty] | mH | mW | microH | milliohm
|  |  |  | 2Z | 0.001 | microH | 1e-06 | 1e-06 | ohm | 1 | 2 | 0.02 | [6 empty] | mH | W | microH | ohm
|  |  |  | 3Z | 1e-05 | milliH | 0.001 | 1e-05 | ohm | 1 | 20 | 0.2 | [6 empty] | mH | W | milliH | ohm
|  |  |  | 4Z | 0.0001 | milliH | 0.001 | 0.0001 | ohm | 1 | 200 | 2 | [6 empty] | mH | ohm | milliH | ohm
|  |  |  | 5Z | 0.001 | milliH | 0.001 | 0.001 | ohm | 1 | 2000 | 20 | [6 empty] | mH | ohm | milliH | ohm
|  |  |  | 6Z | 1e-05 | H | 1 | 1e-05 | kilohm | 1000 | 20000 | 200 | [6 empty] | H | kW | H | kilohm
|  |  |  | 7Z | 0.0001 | H | 1 | 0.0001 | kilohm | 1000 | 200000 | 2000 | [6 empty] | H | kW | H | kilohm
|  |  |  | 1Y | 0.0001 | microF | 1e-06 | 0.0001 | milliS | 0.001 | [8 empty] | mF | mS | microF | milliS
|  |  |  | 2Y | 1e-05 | microF | 1e-06 | 1e-05 | milliS | 0.001 | [8 empty] | mF | mS | microF | milliS
|  |  |  | 3Y | 1e-06 | microF | 1e-06 | 1e-06 | milliS | 0.001 | [8 empty] | mF | mS | microF | milliS
|  |  |  | 4Y | 0.0001 | nF | 1e-09 | 0.0001 | microS | 1e-06 | [8 empty] | nF | mS | nF | microS
|  |  |  | 5Y | 1e-05 | nF | 1e-09 | 1e-05 | microS | 1e-06 | [8 empty] | nF | mS | nF | microS
|  |  |  | 6Y | 0.001 | pF | 1e-12 | 0.001 | nS | 1e-09 | [8 empty] | pF | nS | pF | nS
|  |  |  | 7Y | 0.0001 | pF | 1e-12 | 0.0001 | nS | 1e-09 | [8 empty] | pF | nS | pF | nS
Note on M9: Andrew Corney: capacitor remeasured - use same resistor values as for 22/01/04
Note on N9: Andrew Corney: capacitor remeasured - use same resistor values as for 22/01/04
Note on O9: Andrew Corney: capacitor remeasured - use same resistor values as for 22/01/04
Note on P9: Andrew Corney: capacitor remeasured - use same resistor values as for 22/01/04
Note on Q9: Andrew Corney: capacitor remeasured - use same resistor values as for 22/01/04
Note on R9: Andrew Corney: capacitor remeasured - use same resistor values as for 22/01/04
Note on V18: k.jones: AH2700A check of 1 M at 1.5915 k. AI VII, p.42
Note on M41: k.jones: The n/f or m/f factor for CMC calculation.
Note on L47: k.jones: was just 1.0 in source template

## Uncertainty
UNCERTAINTY |  |  |  |  | Information for calculating expanded uncertainties | [7 empty] | Parameters dependent on component | [6 empty] | This is a data table with B9 as the column input cell. | [13 empty] | CMC |  |  |  | IANZ
GEUM uncertainty calculation template |  |  |  |  | k-factor from degrees of freedom | [14 empty] | This must be re-calculated with new data sets. | [13 empty] | C uncert = 0.2/f+22C |  | delta uncert = 0.000027+0.00027/C
The component label is the key variable for the table |  |  |  |  | a | 1.96 |  |  |  |  | Model for uncertainty calculation. |  | Label | R alpha | u_R alpha | L/C alpha | u_L/C alpha | Aux L | Aux R | Label | L/C uncertainty | k | R/G uncertainty | k | G to delta | Range | L/C | mult | n | R/G | mult | m | f | pF | CMC/U |  | CMC/U | C |  | delta
[5 empty] | b | 2.372 |  |  |  |  | Base on formula for Universal bridge....ie assign an uncertainty to each term. | [9 empty] | 4.2879e-05 | 2.2941 | 0.0003453 | 2.15 | mrad
The 95% level of confidence expanded uncertainties are calculated using the Welch-Satterthwaite |  |  |  |  | c | 1.2737 | [6 empty] | 1 microF3Y | 0 | 0 | 7.5 | 27.5 | 0 | 0 | 1 microF3Y | 4.2879e-05 | 2.2941 | 0.0003453 | 2.15 | 0.034533 | 3Y | 0.99992 | 1e-06 | 0 | 0.0017179 | 0.001 | 0 | 1591.5 | 21.9984 | 0.51303 | 2.7e-05 | 0.78186 | 21.9984 | 0.51303 | 0.03 | 0.86873
formula to combine standard uncertainties. The individual degrees of freedom are derived either |  |  |  |  | d | 7.0962 | [6 empty] | 0.1 microF4Y | 0 | 0 | 7.5 | 27.5 | 0 | 0 | 0.1 microF4Y | 0.0042884 | 2.2941 | 0.034534 | 2.15 | 0.034533 | 4Y | 100.0039 | 1e-09 | 0 | 0.041298 | 1e-06 | 0 | 1591.5 | 2.2002 | 0.51306 | 2.7003e-05 | 0.78193 | 2.2002 | 0.51306 | 0.03 | 0.86873
from the number of measurements used to derive the uncertainty or an estimate | [12 empty] | 0.01 microF5Y | 0 | 0 | 7.5 | 27.5 | 0 | 0 | 0.01 microF5Y | 0.00042887 | 2.2941 | 0.0034536 | 2.15 | 0.034533 | 5Y | 10.001 | 1e-09 | 0 | 0.0049597 | 1e-06 | 0 | 1591.5 | 0.22015 | 0.51332 | 2.7027e-05 | 0.78264 | 0.22015 | 0.51332 | 0.03 | 0.86873
of the percentage uncertainty in the uncertainty. | [12 empty] | 1000 pF6Y | 0 | 0 | 30 | 9 | 0 | 0 | 1000 pF6Y | 0.050273 | 2.1371 | 0.34526 | 2.15 | 0.034531 | 6Y | 999.8688 | 1e-12 | 0 | -0.201 | 1e-09 | 0 | 1591.5 | 0.022123 | 0.44006 | 2.727e-05 | 0.78972 | 0.022123 | 0.44006 | 0.03 | 0.86878
Component label | 1 microF3Y | [11 empty] | 100 pF7Y | 0 | 0 | 30 | 9 | 0 | 0 | 100 pF7Y | 0.0050283 | 2.1371 | 0.034533 | 2.15 | 0.034534 | 7Y | 100.0003 | 1e-12 | 0 | 0.19769 | 1e-09 | 0 | 1591.5 | 0.0023257 | 0.46251 | 2.97e-05 | 0.86002 | 0.0023257 | 0.46251 | 0.03 | 0.86871
Universal bridge reactance uncertainty |  |  |  |  | k=a+b/v+c/v^2+d/v^3 | [7 empty] | 10 pF7Y | 0 | 0 | 30 | 9 | 0 | 0 | 10 pF7Y | 0.00059597 | 2.0955 | 0.0044014 | 2.0613 | 0.044021 | 7Y | 9.9987 | 1e-12 | 0 | 0.019099 | 1e-09 | 0 | 1591.5 | 0.00034564 | 0.57996 | 5.4003e-05 | 1.2268 | 0.00034564 | 0.57996 | 0.03 | 0.6815
| Standard | for | 999846 | L/C dial setting | [8 empty] | 1 pF7Y | 0 | 0 | 40 | 12 | 0 | 0 | 1 pF7Y | 0.00041578 | 2.5532 | 0.0031066 | 2.1827 | 0.31071 | 7Y | 0.99986 | 1e-12 | 0 | 0.0026998 | 1e-09 | 0 | 1591.5 | 0.00014766 | 0.35515 | 0.00029704 | 0.95599 | 0.00014766 | 0.35515 | 0.03 | 0.096553
Source of uncertainty | Uncertainties |  |  | Degrees of | [8 empty] | 100 pF6Y | 0 | 0 | 7.5 | 27.5 | 0 | 0 | 100 pF6Y | 0.0053233 | 2.1647 | 0.044014 | 2.0613 | 0.044014 | 6Y | 100.0038 | 1e-12 | 0 | -0.029 | 1e-09 | 0 | 1591.5 | 0.0023258 | 0.4369 | 2.97e-05 | 0.67478 | 0.0023258 | 0.4369 | 0.03 | 0.6816
| u (counts) | Variances | % Accuracy | Freedom, v | u^4/v |  |  |  |  | Latest comment | [10 empty] | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 0.03 | #N/A
ac/dc error of standards | 2.4996 | 6.2481 | 30 | 5.5556 | 7.0269 | 5ppm considered two sigma limit, essentially relying on dc values of resistors | [14 empty] | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 0.03 | #N/A
dial linearity | 1 | 1 | 30 | 5.5556 | 0.18 | 1 ppm of full scale | [14 empty] | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 0.03 | #N/A
amplifier/ratio | 16.9974 | 288.911 | 30 | 5.5556 | 15024.5213 | based on scatter of results in 1984, increased in 2004 and 2005 |  |  |  | Previous value was 7 ppm, now increased to 17 ppm to acount for errors on RHS of bridge. | [10 empty] | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 0.03 | #N/A
temperature coefficient of resistors | 1.9997 | 3.9988 | 50 | 2 | 7.9951 | depends on matching of Vishay components | [14 empty] | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 0.03 | #N/A
temperature coefficient of capacitor | 0.99985 | 0.99969 | 30 | 5.5556 | 0.17989 | 2 ppm per degree for capacitor | [14 empty] | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 0.03 | #N/A
component stability | 4.9992 | 24.9923 | 50 | 2 | 312.3075 | some doubt on stability of capacitor in particular, increased in 2004 |  |  |  | Increase from 2 to 5 ppm to account for short-term instability of 1000 pF capacitor | [10 empty] | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 0.03 | #N/A
L/C error due to phase error in R/G | 0.0027456 | 7.5383e-06 | - | 13 | 4.3713e-12 | 16 ppm of resistance dial setting, see comment for resistance, reduced to 5 ppm after 2004 cal |  |  |  | Link to conductance setting, previously just 2 counts | [10 empty] | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 0.03 | #N/A
resolution | 1.2567 | 1.5792 | 50 | 2 | 1.247 | this might be range and frequency dependent |  |  |  | Previous was 2, use latest value from MSLT.E.023.005, 200/f | [10 empty] | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 0.03 | #N/A
| 0 | 0 | 1 | 5000 | 0 | [15 empty] | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 0.03 | #N/A
| Sum of variance | 327.7291 |  | Sum of u^4/v | 15353.4577 | [15 empty] | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 0.03 | #N/A
| u combined (counts) | 18.1033 |  | effective v | 6.9956 | [15 empty] | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 0.03 | #N/A
|  |  |  | k | 2.3458 | [15 empty] | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 0.03 | #N/A
Assumptions |  |  |  | expanded u (counts) | 42.4672 | [15 empty] | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 0.03 | #N/A
1/ For all frequencies. Additional noise typical at 50 Hz needs to be added | [20 empty] | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 0.03 | #N/A
2/ Phase angle assumed small, ie no dial interaction. | [20 empty] | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 0.03 | #N/A
3/ Consistent with 1984 calibration, the IEEE paper and recent complement checks | [20 empty] | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 0.03 | #N/A
4/ Updated following 2004/5 UB calibration, see AUDIREGC | [20 empty] | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 0.03 | #N/A
[21 empty] | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 0.03 | #N/A
[21 empty] | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 0.03 | #N/A
Universal bridge resistance uncertainty |  |  |  |  | k=a+b/v+c/v^2+d/v^3 | [15 empty] | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 0.03 | #N/A
| Standard | for | 1716 | conductance dial setting |  | all references to the conductance dial are now to cell D34 (previously separate lookups back to results page) | [14 empty] | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 0.03 | #N/A
Source of uncertainty | Uncertainties |  |  | Degrees of | [16 empty] | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 0.03 | #N/A
| u (counts) | Variances | % Accuracy | Freedom, v | u^4/v |  |  |  |  | Latest comment | [10 empty] | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 0.03 | #N/A
G/R error due to phase error in L/C | 159.9754 | 25592.1158 | - | 13 | 50381260.8835 | based on standard deviation of 1984 determinations, 160 counts (16 ppm) for C at full scale |  |  |  | Previous was 16, use latest value of 5 from MSLT.E.023.005 | Now return to 16 to stay consistent with CMCS | [9 empty] | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 0.03 | #N/A
dial linearity | 10 | 100 | 30 | 5.5556 | 1800 |  |  |  |  | Previous value was 2, now decreased to 1 ppm following MSLT.E023.005 | [10 empty] | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 0.03 | #N/A
ac/dc error of standards | 0.0429 | 0.0018404 | 30 | 5.5556 | 6.0968e-07 | 5ppm considered two sigma limit, essentially relying on dc values of resistors | [14 empty] | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 0.03 | #N/A
resolution | 10 | 100 | 30 | 5.5556 | 1800 | this might be dependent on frequency and range, included below, increased from 0 to 1 in 2006, now 2000/f |  |  |  |  | Increase to 5 for 2011 measurements. | [9 empty] | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 0.03 | #N/A
amplifier/ratio | 0.029172 | 0.00085101 | 30 | 5.5556 | 1.3036e-07 | based on scatter of results in 1984, increased following 2004 and 2005 calibration |  |  |  | Previous value was 7 ppm, now increased to 17 ppm to acount for errors on RHS of bridge. | [10 empty] | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 0.03 | #N/A
temperature coefficient of resistors | 0.03432 | 0.0011779 | 50 | 2 | 6.9368e-07 | depends on matching of Vishay components | [14 empty] | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 0.03 | #N/A
component stability | 0.03432 | 0.0011779 | 30 | 5.5556 | 2.4972e-07 | stability of resistors only | [14 empty] | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 0.03 | #N/A
| 0 | 0 | 50 | 2 | 0 |  |  |  |  | Previous value was 1, now increased to 2 ppm following MSLT.E023.005 | [10 empty] | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 0.03 | #N/A
| 0 | 0 | 1 | 5000 | 0 | [15 empty] | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 0.03 | #N/A
| Sum of variance | 25792.1209 |  | Sum of u^4/v | 50384860.8835 | [15 empty] | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 0.03 | #N/A
| u combined (counts) | 160.5993 |  | effective v | 13.203 | [15 empty] | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 0.03 | #N/A
|  |  |  | k | 2.15 | [15 empty] | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 0.03 | #N/A
Assumptions |  |  |  | expanded u (counts) | 345.2957 | [15 empty] | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 0.03 | #N/A
1/ This scatter is consistent with Thompson's phase standards. | [20 empty] | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 0.03 | #N/A
2/ Updated following 200/54 UB calibration see AUDIREGC | [20 empty] | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 0.03 | #N/A
[21 empty] | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 0.03 | #N/A
[21 empty] | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 0.03 | #N/A
Conductance uncertainty for Muirhead capacitors |  |  |  |  | k=a+b/v+c/v^2+d/v^3 | [15 empty] | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 0.03 | #N/A
| Standard | [19 empty] | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 0.03 | #N/A
Source of uncertainty | Uncertainties |  |  | Degrees of | [21 empty] | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A
| u (counts) | Variances | % Accuracy | Freedom, v | u^4/v | [20 empty] | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A
ubridge uncertainty | 160.5993 | 25792.1209 | - | 13.203 | 50384860.8835 | ubridge count uncertainty | [19 empty] | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A
3 month drift in resistance | 0 | 0 | 30 | 5.5556 | 0 | [20 empty] | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A
temperature coefficient | 0 | 0 | 20 | 12.5 | 0 | assume u_t = 0.5 degree
series resistance | 0 | 0 | 30 | 5.5556 | 0 | a series R will map to a conductance, worst case 10 milliohm in series, G = (wC)^2*R | [27 empty] | C uncert = (0.2/f +22C) pF with C in microF, f in Hz
| Sum of variance | 25792.1209 |  | Sum of u^4/v | 50384860.8835
| u combined (counts) | 160.5993 |  | effective v | 13.203
|  |  |  | k | 2.15 |  | range | cap | Cmult | G | Gmult | unit mult
Assumptions |  |  |  | expanded u (counts) | 345.2957 |  | 3Y | 0.99992 | 1e-06 | 0.0017179 | 1e-06 | 0.001
1/ |  |  |  |  | 0.0003453 | milliS
Reactance uncertainty for Muirhead capacitors |  |  |  |  | k=a+b/v+c/v^2+d/v^3
Standard
Source of uncertainty | Uncertainties |  |  | Degrees of
| u (counts) | Variances | % Accuracy | Freedom, v | u^4/v
ubridge uncertainty | 18.1033 | 327.7291 | - | 6.9956 | 15353.4577
temperature coefficient | 4.6496 | 21.6183 | 30 | 8 | 58.4191 | assume 0.5 deg at one sigma and look up the temp coefficient, 50 ppm for the temp coefficient (ref. Kibble &Rayner p. 61)
3 month drift of standard | 0 | 0 | 30 | 5.5556 | 0
Connection definition | 0 | 0 | 30 | 5.5556 | 0 | 0.1 pF for all capacitors at 2 sigma level, see AI V 1/11/95 | [5 empty] | set to zero for 4 terminal-pair definition
| Sum of variance | 349.3474 |  | Sum of u^4/v | 15411.8768
| u combined (count) | 18.6908 |  | effective v | 7.9188
|  |  |  | k | 2.2941
Assumptions |  |  |  | expanded u (count) | 42.8794
1/ |  |  |  |  | 4.2879e-05 | microF
0.99992
Note on W8: k.jones: 30% of HP quoted TC for uncertainty
Note on B18: t.stewart: Assume 0.5 degree as have below? KJ: agreed and modified
Note on B20: t.stewart: 5 or 16 ppm?? Latest comment and MSLT.E023.005 suggests 5 ppm. KJ: agree 5 in 2011 but in 2014, safer to stay with 16
Note on B21: k.jones: Sensitivity falls with f, 1 ppm at 2 kHz
Note on B37: t.stewart: 5 or 16 ppm?? Latest comment and MSLT.E023.005 suggests 5 ppm. KJ: agree 5 in 2011 but revert to 16 in 2014
Note on L37: k.jones: Revert after new CMCs developed and accepted.
Note on D38: k.jones: changed from 50 to 30 %, has impact on dof.
Note on B40: k.jones: say 1/f noise of 1 ppm at 100 Hz
Note on D40: k.jones: Was set to 1%, 30% compatible with CMC calculation
Note on L40: k.jones: 10 in last dial is more consistent with CMCs.
Note on B41: k.jones: was with x10 but!!! the dial setting takes care of the fact that there are 7 dials!!!!
Note on B57: k.jones: Was done in % for the Muirhead inductors.
Note on B59: k.jones: set to zero until rest of calculation is correct.
Note on B60: k.jones: 2nd order formula is rss of alpha*u_t and u_t*u_alpha.
Note on B61: Keith Jones: series R not an issue with four term pair definition
Note on B63: t.stewart: Should it be "u combined (%)" ?? KJ: agree %
Note on B72: k.jones: Was % for Muirheads
Note on B73: k.jones: 2nd order formula is rss of alpha*u_t and u_t*u_alpha.
Note on B74: k.jones: set to zero for now
Note on B75: Keith Jones: 1 pF is 1000 counts on range 6Y.

## Report
Formatted version | [17 empty] | 2/3/2010 results |  | S21759
[8 empty] | dial code | Serial | Nominal | Capacitance | Expanded | k | Loss angle | Expanded | k | Difference as En |  | Serial | Nominal | Capacitance | Expanded | k | Loss Angle | Expanded | k
[9 empty] | Number | Value |  | uncertainty |  |  | uncertainty |  |  |  | Number | Value |  | Uncertainty |  |  | Uncertainty
Serial Number | Nominal Value (µF) | Capacitance (nF) | Expanded Uncertainty (nF) | k | Loss Angle (mrad) | Expanded Uncertainty (mrad) | k | [6 empty] | (milliradian) | (milliradian) | [8 empty] | (milliradian) | (milliradian)
0 | 1 | 999868.8292 | 50.2726 | 2.1371 | -0.020103 | 0.034531 | 2.15 | 1000 pF6Y |  | 1000 | 999.8688 | 0.050273 | 2.1371 | -0.020103 | 0.034531 | 2.15 | 19869.0531 | -2.4532 | 365442 | 1.0 mF | 0.99984 | 6e-05 | 2.0825 | 0.064607 | 0.12412 | 2.3634
0 | 0.1 | 100.0003 | 0.0050283 | 2.1371 | 0.19769 | 0.034534 | 2.15 | 100 pF7Y |  | 100 | 100.0003 | 0.0050283 | 2.1371 | 0.19769 | 0.034534 | 2.15 | -19.2339 | 4.041 | 365383 | 0.1 mF | 100.097 | 0.006 | 2.0825 | 0.058141 | 0.035704 | 2.1186
0 | 0.01 | 9.9987 | 0.00059597 | 2.0955 | 0.19102 | 0.044021 | 2.0613 | 10 pF7Y |  | 10 | 9.9987 | 0.00059597 | 2.0955 | 0.19102 | 0.044021 | 2.0613 | 11.122 | 2.4365 | 365381 | 0.01 mF | 9.9921 | 0.0007 | 2.0763 | 0.083762 | 0.034588 | 2.1481
0 | 0.001 | 0.00099986 | 4.1578e-07 | 2.5532 | 0.27003 | 0.31071 | 2.1827 | 1 pF7Y |  | 1 | 0.99986 | 0.00041578 | 2.5532 | 0.27003 | 0.31071 | 2.1827 | -2374100.2235 | 0.79056 | 365391 | 0.001 mF | 988.11 | 0.12 | 2.0724 | 0.02439 | 0.034607 | 2.1484
0 | 0.0005 | 0.00099992 | 4.2879e-08 | 2.2941 | 0.17181 | 0.034533 | 2.15 | 1 microF3Y |  | 1 | 0.99992 | 4.2879e-05 | 2.2941 | 0.17181 | 0.034533 | 2.15 | -11416893.9742 | 3.0212 | 365396 | 0.0005 mF | 490.55 | 0.11 | 2.1266 | 0.067476 | 0.035041 | 2.1352
[8 empty] | 0.1 microF4Y |  | 0.1 | 0.1 | 4.2884e-06 | 2.2941 | 0.041297 | 0.034533 | 2.15
[8 empty] | 0.01 microF5Y |  | 0.01 | 0.010001 | 4.2887e-07 | 2.2941 | 0.049594 | 0.034533 | 2.15
[8 empty] | 100 pF6Y |  | 100 | 100.0038 | 0.0053233 | 2.1647 | -0.029 | 0.044014 | 2.0613

## Conditions
| Temp |  | RH
All GMH5I readings | 20.77 | 0.091378
Room | 21.4 | 21.3 | 50 | 50.1

## RoomLogger
2017-10-26 08:45 | 21.4 | 49.4
2017-10-26 08:50 | 21.4 | 49.3
2017-10-26 08:55 | 21.4 | 49.3
2017-10-26 09:00 | 21.4 | 49.4
2017-10-26 09:05 | 21.4 | 49.5
2017-10-26 09:10 | 21.4 | 49.7
2017-10-26 09:15 | 21.3 | 49.9
2017-10-26 09:20 | 21.4 | 49.7
2017-10-26 09:25 | 21.4 | 49.6
2017-10-26 09:30 | 21.4 | 49.5
2017-10-26 09:35 | 21.4 | 49.8
2017-10-26 09:40 | 21.4 | 49.9
2017-10-26 09:45 | 21.3 | 50.1
2017-10-26 09:50 | 21.3 | 50
2017-10-26 09:55 | 21.4 | 49.8
2017-10-26 10:00 | 21.4 | 49.8
2017-10-26 10:05 | 21.4 | 49.8
2017-10-26 10:10 | 21.4 | 49.7
2017-10-26 10:15 | 21.4 | 49.7
2017-10-26 10:20 | 21.4 | 49.7
2017-10-26 10:25 | 21.4 | 49.8
2017-10-26 10:30 | 21.4 | 49.6
2017-10-26 10:35 | 21.4 | 49.6
2017-10-26 10:40 | 21.3 | 49.8 | 21.3125 | 50.1062
2017-10-26 10:45 | 21.4 | 49.8 | 0.034157 | 0.1652
2017-10-26 10:50 | 21.3 | 50.1
2017-10-26 10:55 | 21.3 | 50.1
2017-10-26 11:00 | 21.3 | 50.1
2017-10-26 11:05 | 21.3 | 50
2017-10-26 11:10 | 21.4 | 50
2017-10-26 11:15 | 21.3 | 50.1
2017-10-26 11:20 | 21.3 | 50
2017-10-26 11:25 | 21.3 | 50.2
2017-10-26 11:30 | 21.3 | 50.2
2017-10-26 11:35 | 21.3 | 50.1
2017-10-26 11:40 | 21.3 | 50.3
2017-10-26 11:45 | 21.3 | 50.3
2017-10-26 11:50 | 21.3 | 50.2
2017-10-26 11:55 | 21.3 | 50.4
2017-10-26 12:00 | 21.3 | 50.3
2017-10-26 12:05 | 21.3 | 50.1
2017-10-26 12:10 | 21.3 | 49.9
2017-10-26 12:15 | 21.3 | 49.9
2017-10-26 12:20 | 21.3 | 49.8
2017-10-26 12:25 | 21.4 | 49.8
2017-10-26 12:30 | 21.3 | 49.8
2017-10-26 12:35 | 21.3 | 49.8
2017-10-26 12:40 | 21.3 | 49.8
2017-10-26 12:45 | 21.4 | 49.7
2017-10-26 12:50 | 21.4 | 49.7
2017-10-26 12:55 | 21.3 | 49.8
2017-10-26 13:00 | 21.3 | 49.7
2017-10-26 13:05 | 21.3 | 49.7
2017-10-26 13:10 | 21.3 | 49.9
2017-10-26 13:15 | 21.3 | 49.8
2017-10-26 13:20 | 21.3 | 49.9
2017-10-26 13:25 | 21.3 | 50.1
2017-10-26 13:30 | 21.3 | 50.1
2017-10-26 13:35 | 21.3 | 50.2
2017-10-26 13:40 | 21.4 | 50
2017-10-26 13:45 | 21.4 | 49.9
2017-10-26 13:50 | 21.4 | 50
2017-10-26 13:55 | 21.3 | 50.1
2017-10-26 14:00 | 21.4 | 50.2
2017-10-26 14:05 | 21.4 | 50
2017-10-26 14:10 | 21.4 | 49.8
2017-10-26 14:15 | 21.4 | 49.7
2017-10-26 14:20 | 21.4 | 49.6
2017-10-26 14:25 | 21.3 | 49.6
2017-10-26 14:30 | 21.4 | 49.5
2017-10-26 14:35 | 21.3 | 49.5
2017-10-26 14:40 | 21.3 | 49.6
2017-10-26 14:45 | 21.3 | 49.7
2017-10-26 14:50 | 21.4 | 49.7
2017-10-26 14:55 | 21.4 | 49.7
2017-10-26 15:00 | 21.3 | 49.7
2017-10-26 15:05 | 21.3 | 49.7
2017-10-26 15:10 | 21.3 | 49.6
2017-10-26 15:15 | 21.3 | 49.6
2017-10-26 15:20 | 21.3 | 49.7
2017-10-26 15:25 | 21.3 | 49.6
2017-10-26 15:30 | 21.3 | 49.6
2017-10-26 15:35 | 21.2 | 49.8
2017-10-26 15:40 | 21.2 | 49.8
2017-10-26 15:45 | 21.3 | 49.7
2017-10-26 15:50 | 21.3 | 49.6
2017-10-26 15:55 | 21.3 | 49.6
2017-10-26 16:00 | 21.3 | 49.8
2017-10-26 16:05 | 21.3 | 49.8
2017-10-26 16:10 | 21.3 | 49.6
2017-10-26 16:15 | 21.3 | 49.6
2017-10-26 16:20 | 21.3 | 49.7
2017-10-26 16:25 | 21.3 | 49.6
2017-10-26 16:30 | 21.3 | 49.6
2017-10-26 16:35 | 21.3 | 49.4
2017-10-26 16:40 | 21.3 | 49.5
2017-10-26 16:45 | 21.3 | 49.4
2017-10-26 16:50 | 21.3 | 49.5
2017-10-26 16:55 | 21.3 | 49.4
2017-10-26 17:00 | 21.3 | 49.4
2017-10-26 17:05 | 21.3 | 49.5
2017-10-26 17:10 | 21.3 | 49.5
2017-10-26 17:15 | 21.3 | 49.5
2017-10-26 17:20 | 21.3 | 49.3
2017-10-26 17:25 | 21.3 | 49.3
2017-10-26 17:30 | 21.2 | 49.4
2017-10-26 17:35 | 21.3 | 49.4
2017-10-26 17:40 | 21.3 | 49.4
2017-10-26 17:45 | 21.3 | 49.3
2017-10-26 17:50 | 21.3 | 49.3
2017-10-26 17:55 | 21.3 | 49.2
2017-10-26 18:00 | 21.3 | 49.2
2017-10-26 18:05 | 21.3 | 49.2
2017-10-26 18:10 | 21.3 | 49.3
2017-10-26 18:15 | 21.3 | 49.2
2017-10-26 18:20 | 21.3 | 49.2
2017-10-26 18:25 | 21.3 | 49.3
2017-10-26 18:30 | 21.3 | 49.4
2017-10-26 18:35 | 21.3 | 49.4
2017-10-26 18:40 | 21.3 | 49.5
2017-10-26 18:45 | 21.2 | 49.5
2017-10-26 18:50 | 21.3 | 49.5
2017-10-26 18:55 | 21.3 | 49.5
2017-10-26 19:00 | 21.3 | 49.6
2017-10-26 19:05 | 21.3 | 49.6
2017-10-26 19:10 | 21.3 | 49.6
2017-10-26 19:15 | 21.3 | 49.7
2017-10-26 19:20 | 21.3 | 49.8
2017-10-26 19:25 | 21.3 | 49.8
2017-10-26 19:30 | 21.3 | 49.8
2017-10-26 19:35 | 21.3 | 49.7
2017-10-26 19:40 | 21.3 | 49.7
2017-10-26 19:45 | 21.3 | 49.6
2017-10-26 19:50 | 21.3 | 49.8
2017-10-26 19:55 | 21.3 | 49.7
2017-10-26 20:00 | 21.4 | 49.6
2017-10-26 20:05 | 21.4 | 49.7
2017-10-26 20:10 | 21.3 | 49.8
2017-10-26 20:15 | 21.4 | 49.8
2017-10-26 20:20 | 21.3 | 49.9
2017-10-26 20:25 | 21.4 | 49.7
2017-10-26 20:30 | 21.3 | 49.9
2017-10-26 20:35 | 21.4 | 49.8
2017-10-26 20:40 | 21.4 | 49.8
2017-10-26 20:45 | 21.3 | 50
2017-10-26 20:50 | 21.4 | 49.7
2017-10-26 20:55 | 21.3 | 49.8
2017-10-26 21:00 | 21.3 | 49.8
2017-10-26 21:05 | 21.3 | 49.8
2017-10-26 21:10 | 21.3 | 49.9
2017-10-26 21:15 | 21.4 | 49.7
2017-10-26 21:20 | 21.3 | 49.7
2017-10-26 21:25 | 21.3 | 49.7
2017-10-26 21:30 | 21.3 | 49.7
2017-10-26 21:35 | 21.3 | 49.7
2017-10-26 21:40 | 21.4 | 49.7
2017-10-26 21:45 | 21.3 | 49.7
2017-10-26 21:50 | 21.4 | 49.7
2017-10-26 21:55 | 21.4 | 49.6
2017-10-26 22:00 | 21.4 | 49.7
2017-10-26 22:05 | 21.4 | 49.7
2017-10-26 22:10 | 21.4 | 49.7
2017-10-26 22:15 | 21.3 | 49.7
2017-10-26 22:20 | 21.3 | 49.7
2017-10-26 22:25 | 21.4 | 49.6
2017-10-26 22:30 | 21.3 | 49.7
2017-10-26 22:35 | 21.4 | 49.5
2017-10-26 22:40 | 21.4 | 49.5
2017-10-26 22:45 | 21.4 | 49.4
2017-10-26 22:50 | 21.4 | 49.4
2017-10-26 22:55 | 21.3 | 49.4
2017-10-26 23:00 | 21.3 | 49.4
2017-10-26 23:05 | 21.3 | 49.4
2017-10-26 23:10 | 21.4 | 49.3
2017-10-26 23:15 | 21.4 | 49.4
2017-10-26 23:20 | 21.3 | 49.5
2017-10-26 23:25 | 21.3 | 49.5
2017-10-26 23:30 | 21.4 | 49.3
2017-10-26 23:35 | 21.3 | 49.5
2017-10-26 23:40 | 21.4 | 49.4
2017-10-26 23:45 | 21.3 | 49.5
2017-10-26 23:50 | 21.4 | 49.3
2017-10-26 23:55 | 21.3 | 49.4
2017-10-27 | 21.4 | 49.3
2017-10-27 00:05 | 21.4 | 49.3
2017-10-27 00:10 | 21.4 | 49.2
2017-10-27 00:15 | 21.4 | 49.2
2017-10-27 00:20 | 21.4 | 49.2
2017-10-27 00:25 | 21.4 | 49.2
2017-10-27 00:30 | 21.4 | 49.3
2017-10-27 00:35 | 21.4 | 49.3
2017-10-27 00:40 | 21.4 | 49.2
2017-10-27 00:45 | 21.3 | 49.4
2017-10-27 00:50 | 21.4 | 49.4
2017-10-27 00:55 | 21.4 | 49.4
2017-10-27 01:00 | 21.4 | 49.2
2017-10-27 01:05 | 21.4 | 49.2
2017-10-27 01:10 | 21.4 | 49.2
2017-10-27 01:15 | 21.4 | 49.2
2017-10-27 01:20 | 21.4 | 49.2
2017-10-27 01:25 | 21.4 | 49.2
2017-10-27 01:30 | 21.4 | 49.2
2017-10-27 01:35 | 21.3 | 49.3
2017-10-27 01:40 | 21.3 | 49.3
2017-10-27 01:45 | 21.4 | 49.1
2017-10-27 01:50 | 21.3 | 49.2
2017-10-27 01:55 | 21.4 | 49.1
2017-10-27 02:00 | 21.4 | 49.2
2017-10-27 02:05 | 21.4 | 49.1
2017-10-27 02:10 | 21.4 | 49.1
2017-10-27 02:15 | 21.4 | 49.2
2017-10-27 02:20 | 21.4 | 49
2017-10-27 02:25 | 21.4 | 49.1
2017-10-27 02:30 | 21.4 | 49.1
2017-10-27 02:35 | 21.3 | 49.1
2017-10-27 02:40 | 21.4 | 49.1
2017-10-27 02:45 | 21.4 | 49
2017-10-27 02:50 | 21.4 | 49
2017-10-27 02:55 | 21.4 | 49
2017-10-27 03:00 | 21.4 | 49
2017-10-27 03:05 | 21.4 | 49.1
2017-10-27 03:10 | 21.4 | 48.9
2017-10-27 03:15 | 21.4 | 48.8
2017-10-27 03:20 | 21.4 | 49
2017-10-27 03:25 | 21.4 | 48.9
2017-10-27 03:30 | 21.4 | 49
2017-10-27 03:35 | 21.4 | 48.9
2017-10-27 03:40 | 21.4 | 48.9
2017-10-27 03:45 | 21.3 | 49
2017-10-27 03:50 | 21.4 | 48.8
2017-10-27 03:55 | 21.4 | 48.8
2017-10-27 04:00 | 21.4 | 48.7
2017-10-27 04:05 | 21.4 | 48.7
2017-10-27 04:10 | 21.4 | 48.7
2017-10-27 04:15 | 21.4 | 48.6
2017-10-27 04:20 | 21.4 | 48.6
2017-10-27 04:25 | 21.4 | 48.6
2017-10-27 04:30 | 21.4 | 48.5
2017-10-27 04:35 | 21.4 | 48.5
2017-10-27 04:40 | 21.4 | 48.5
2017-10-27 04:45 | 21.4 | 48.5
2017-10-27 04:50 | 21.4 | 48.4
2017-10-27 04:55 | 21.4 | 48.4
2017-10-27 05:00 | 21.4 | 48.3
2017-10-27 05:05 | 21.4 | 48.3
2017-10-27 05:10 | 21.5 | 48.2
2017-10-27 05:15 | 21.4 | 48.3
2017-10-27 05:20 | 21.4 | 48.2
2017-10-27 05:25 | 21.4 | 48.1
2017-10-27 05:30 | 21.4 | 48.1
2017-10-27 05:35 | 21.4 | 48
2017-10-27 05:40 | 21.4 | 48.2
2017-10-27 05:45 | 21.4 | 48
2017-10-27 05:50 | 21.4 | 48
2017-10-27 05:55 | 21.4 | 48
2017-10-27 06:00 | 21.4 | 48
2017-10-27 06:05 | 21.4 | 48
2017-10-27 06:10 | 21.4 | 47.9
2017-10-27 06:15 | 21.4 | 48
2017-10-27 06:20 | 21.4 | 47.8
2017-10-27 06:25 | 21.4 | 47.8
2017-10-27 06:30 | 21.4 | 47.8
2017-10-27 06:35 | 21.4 | 47.8
2017-10-27 06:40 | 21.4 | 47.8
2017-10-27 06:45 | 21.3 | 47.9
2017-10-27 06:50 | 21.4 | 47.7
2017-10-27 06:55 | 21.4 | 47.6
2017-10-27 07:00 | 21.4 | 47.6
2017-10-27 07:05 | 21.4 | 47.6
2017-10-27 07:10 | 21.4 | 47.6
2017-10-27 07:15 | 21.4 | 47.5
2017-10-27 07:20 | 21.5 | 47.4
2017-10-27 07:25 | 21.4 | 47.5
2017-10-27 07:30 | 21.4 | 47.5
2017-10-27 07:35 | 21.4 | 47.4
2017-10-27 07:40 | 21.4 | 47.5
2017-10-27 07:45 | 21.4 | 47.4
2017-10-27 07:50 | 21.4 | 47.4
2017-10-27 07:55 | 21.4 | 47.4
2017-10-27 08:00 | 21.4 | 47.4
2017-10-27 08:05 | 21.4 | 47.4
2017-10-27 08:10 | 21.5 | 47.2
2017-10-27 08:15 | 21.4 | 47.3
2017-10-27 08:20 | 21.4 | 47.3
2017-10-27 08:25 | 21.5 | 47.2
2017-10-27 08:30 | 21.5 | 47.1
2017-10-27 08:35 | 21.4 | 47.2
2017-10-27 08:40 | 21.4 | 47.2

## MuirheadHistory
S21283 | 1998-12-01 |  | Nominal | Capacitance |  | Uncertainty |  | Coverage | Loss Angle | Uncertainty | Coverage
|  |  | Value |  |  |  |  | Factor | (milliradians) | (milliradians) | Factor
|  |  | 1 mF | 0.99931 | mF | 6.1732e-05 | mF | 2.4204 | 0.086757 | 0.034492 | 2.09
|  |  | 100 nF | 100.164 | nF | 0.0061881 | nF | 2.42 | 0.077397 | 0.034494 | 2.1516
|  |  | 10 nF | 9.9944 | nF | 0.00065267 | nF | 2.208 | 0.0115 | 0.039661 | 2.1516
|  |  | 1 nF | 0.98854 | nF | 0.00012821 | nF | 2.2631 | 0.075454 | 0.026289 | 2.0985
|  |  | 500 pF | 490.6498 | pF | 0.1294 | pF | 2.236 | 0.20161 | 0.05114 | 2.2724
S21438 | 2001-12-03 |  | Nominal Value | Capacitance |  | Uncertainty |  | Coverage | Loss Angle | Uncertainty | Coverage
[8 empty] | Factor | (milliradians) | (milliradians) | Factor
|  |  | 1 mF | 0.99983 | mF | 0.0001237 | mF | 2.3983 | 0.052007 | 0.039873 | 2.146
|  |  | 100 nF | 100.0907 | nF | 0.01237 | nF | 2.3995 | 0.06284 | 0.039873 | 2.146
|  |  | 10 nF | 9.9928 | nF | 0.0015145 | nF | 2.1126 | 0.18814 | 0.10942 | 2.146
|  |  | 1 nF | 0.98796 | nF | 0.0001814 | nF | 2.0754 | 0.080975 | 0.10942 | 2.146
|  |  | 500 pF | 490.43 | pF | 0.1347 | pF | 2.1682 | 0.14273 | 0.10942 | 2.146
S21584 | 2005-02-10 |  | Nominal value | Capacitance |  | Expanded uncertainty |  | Coverage factor | Loss angle (milliradians) | Expanded uncertainty (milliradians) | Coverage factor
[8 empty] | Factor |  | uncertainty | Factor
[9 empty] | (milliradian) | (milliradian)
|  |  | 1mF | 0.99934 | mF | 0.00013352 | mF | 2.053 | 0.080723 | 0.03455 | 2.1447
|  |  | 100nF | 100.1763 | nF | 0.017502 | nF | 2.0586 | 0.050503 | 0.034522 | 2.1457
|  |  | 10nF | 9.9949 | nF | 0.001166 | nF | 2.0906 | 0.08091 | 0.034523 | 2.1457
|  |  | 1nF | 0.98942 | nF | 0.0003367 | nF | 2.0793 | 0.042564 | 0.034525 | 2.1456
|  |  | 500pF | 491.6507 | pF | 0.60025 | pF | 2.1249 | 0.13877 | 0.035259 | 2.1286
S22012 | 2017-10-26 | 1000 | 999.8688 | 5.0273e-05 |  | 0.0021371 |  | -0.020103 | 0.034531 | 2.15 | #REF!
|  | 100 | 100.0003 | 0.0050283 |  | 2.1371 |  | 0.19769 | 0.034534 | 2.15 | #REF!
|  | 10 | 9.9987 | 0.00059597 |  | 2.0955 |  | 0.19102 | 0.044021 | 2.0613 | #REF!
|  | 1 | 0.99986 | 0.00041578 |  | 2.5532 |  | 0.27003 | 0.31071 | 2.1827 | #REF!
|  | #REF! | #REF! | #REF! |  | #REF! |  | #REF! | #REF! | #REF! | #REF!
% difference from start
|  | 1000 | 100 | 10 | 1 | #REF! |  |  | 307348 | 365390 | 365385 | 365392 | 365394
| 1998-12-01 | 0.99931 | 100.164 | 9.9944 | 0.98854 | 490.6498 |  | 1998-12-01 | 0 | 0 | 0 | 0 | 0
| 2001-12-03 | 0.99983 | 100.0907 | 9.9928 | 0.98796 | 490.43 |  | 2001-12-03 | 0.052288 | -0.07324 | -0.015524 | -0.058635 | -0.044802
| 2005-02-10 | 0.99934 | 100.1763 | 9.9949 | 0.98942 | 491.6507 |  | 2005-02-10 | 0.0029039 | 0.012278 | 0.005605 | 0.088987 | 0.204
| 2017-10-26 | 5.0273e-05 | 0.0050283 | 0.00059597 | 0.00041578 | #REF! |  | 2017-10-26 | -99.995 | -99.995 | -99.994 | -99.9579 | #REF!

## AH2700Results
| 100 pF dial | [6 empty] | 9999.6894 |  |  |  | zero if AH used for report
|  | .Dial | C | G | t | delta C | delta G |  | report | uncert | En | AH2700-report
|  |  | pF | nS | deg celsius | pF | nS | mrad | mrad | mrad |  | mrad
D106Y |  | 10 | 1034.1413 | 2.1624 | 19.42 | 984.2794 | 0.76009 | 0.077226 | #REF! | #REF! | #REF! | #REF!
D96Y |  | 9 | 937.5681 | 1.8522 | 19.42 | 887.7062 | 0.44994 | 0.050688 | #REF! | #REF! | #REF! | #REF!
D86Y |  | 8 | 837.8403 | 2.0617 | 19.42 | 787.9783 | 0.65942 | 0.083688 | #REF! | #REF! | #REF! | #REF!
D76Y |  | 7 | 741.0266 | 1.7486 | 19.43 | 691.1646 | 0.3463 | 0.050106 | #REF! | #REF! | #REF! | #REF!
D66Y |  | 6 | 641.9902 | 1.9633 | 19.43 | 592.1283 | 0.56098 | 0.094743 | #REF! | #REF! | #REF! | #REF!
D56Y |  | 5 | 539.0917 | 1.9137 | 19.44 | 489.2298 | 0.51139 | 0.10453 | #REF! | #REF! | #REF! | #REF!
D46Y |  | 4 | 442.2966 | 1.5976 | 19.44 | 392.4347 | 0.19531 | 0.04977 | #REF! | #REF! | #REF! | #REF!
D36Y |  | 3 | 342.5975 | 1.8102 | 19.42 | 292.7356 | 0.4079 | 0.13935 | #REF! | #REF! | #REF! | #REF!
D26Y |  | 2 | 245.7763 | 1.4956 | 19.41 | 195.9144 | 0.093346 | 0.047648 | #REF! | #REF! | #REF! | #REF!
D16Y |  | 1 | 146.7929 | 1.7152 | 19.4 | 96.931 | 0.3129 | 0.32282 | #REF! | #REF! | #REF! | #REF!
|  | 0 | 49.8619 | 1.4023 | 19.41 | 0 | 0 |  |  |  |  | #REF!
UB1000 measurement from p.33 of UBIII
|  | pF | nS | mrad
|  | 1000.1648 | 0.4913 | 0.049123 | [8 empty] | Note
The dials each use 4 capacitors with values 1, 2, 2 and 5.
| Note | [11 empty] | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10
| The UB calculation does not assume a 50 urad phase for UB1000 | [11 empty] | 1 | 2 | 1+2 | 2+2 | 1+2+2 | 1+5 | 2+5 | 1+2+5 | 2+2+5 | 1+2+2+5
| Long leads might compromise the AH2700A measurement of UB1000 loss angle. | [11 empty] | The dials each use 4 capacitors with values 1, 2, 2 and 5. 1: All dial settings that include the single 100 pF cap (highlighted in red) have an 'excess' G of about 0.2 nS
So might have once had 0.05 nS for the 100 pF (additional to box zero), but now have 0.3 nS.
[6 empty] | delta C | delta G |  | report | uncert | En | AH2700-report | The dials each use 4 capacitors with values 1, 2, 2 and 5. 1: 4000000000; The dials each use 4 capacitors with values 1, 2, 2 and 5. 2: ohms additional leakage across the capacitor could account for the change.
|  |  | pF | nS | t | pF | nS | mrad | mrad | mrad |  | mrad
| 1pF DIV | 0 | pF: 49.9305; nS: 1.4287; t: 19.4
E16Y |  | 10 | 60.0587 | 1.438 | 19.4 | 10.1281 | 0.00932 | 0.092024 | -0.020103 | 0.034531 | 3.2471 | 0.11213
E26Y |  | 20 | 70.9559 | 1.4565 | 19.39 | 21.0254 | 0.02788 | 0.13261 | 0.19769 | 0.034534 | -1.8847 | -0.065086
E36Y |  | 30 | 81.4923 | 1.4451 | 19.39 | 31.5618 | 0.01648 | 0.052217 | 0.19102 | 0.044021 | -3.1531 | -0.1388
E46Y |  | 40 | 91.6725 | 1.4492 | 19.39 | 41.7419 | 0.02054 | 0.049209 | 0.27003 | 0.31071 | -0.71069 | -0.22082
E56Y |  | 50 | 101.3124 | 1.4483 | 19.38 | 51.3818 | 0.01968 | 0.038303 | 0.17181 | 0.034533 | -3.866 | -0.13351
E66Y |  | 60 | 111.0929 | 1.4488 | 19.4 | 61.1624 | 0.02015 | 0.032946 | #REF! | #REF! | #REF! | #REF!
E76Y |  | 70 | 120.5794 | 1.4629 | 19.4 | 70.6488 | 0.03423 | 0.048452 | #REF! | #REF! | #REF! | #REF!
E86Y |  | 80 | 130.2826 | 1.4672 | 19.41 | 80.3521 | 0.03856 | 0.04799 | #REF! | #REF! | #REF! | #REF!
E96Y |  | 90 | 140.2682 | 1.4662 | 19.41 | 90.3377 | 0.03754 | 0.041556 | #REF! | #REF! | #REF! | #REF!
E106Y |  | 100 | 150.2697 | 1.4538 | 19.41 | 100.3392 | 0.02514 | 0.025056 | #REF! | #REF! | #REF! | #REF!
| 1000 pF step | 1 | 1045.5838 | 1.9795 | 19.43 | 995.6509 | 0.52968 | 0.053201 | #REF! | #REF! | #REF! | #REF!
| 100 pF step | 1 | 146.8762 | 1.7225 | 19.41 | 96.9433 | 0.27269 | 0.2813 | #REF! | #REF! | #REF! | #REF!
E06Y |  | 0 | 49.9329 | 1.4498 | 19.46 | 49.0423 | 1.4497 | 2.9562 | #REF! | #REF! | #REF! | #REF!
| Cu jig |  | pF: 0.89058; nS: 3.61e-05

## Range7
Universal Bridge | Range 7 was a design afterthought on the Universal Bridge and is expected to perform less well than the other ranges.
Consider the 2014 calibration of the UB that has, for the first time, included a 1 M Vishay in the calibration.
1 | At 1.6 kHz the G and R ac-dc differences look OK at -10.5 and 0.3 ppm respectively (using the same 'c' correction as for all other ranges, see 'Resistors' page).
We have no resistance phase standard for range 7.
2 | The 1 M Vishay gives a reading of -35 mH, which is much larger than the intrinsic <0.08 uH, but consistent with a mere 0.035 pF in parallel with the 1 M resistor
3 | GR100 is measured ('L and C' sheet) to have a 144 urad loss angle.
4 | AH2700 measures |  | 5.353 | urad for GR100 (see p.23 of UB III)
5 | Calc cap suggests |  | -0.02 | urad for GR100 (see Y:\Ongoing\Farad\CalcCap\Cap2014\Simple_buildup_July2014.xls)
Item 1 gives basic confidence of ratio performance.
Item 2 is inconclusive.
Items 3 to 5 suggest a correction is required for phase on Range 7Y of the UB.
Unfortunately the difference in loss angles measured by the UB in July and August exceed the estimated uncertainty.
It is possible that the phase of the UB's 1 M internal standard was changed during the July calibration (physical knock?)
AH2700 | Therefore it seems prudent to rely on the AH2700 for the loss angle of capacitors below 1 nF.
See AH2700 manual page C-6 (appendix C), 'Accuracy of D vs. Frequency' graph
D accuracy
| 1 pF | 6 urad
| 10 pF | 4 urad
| 100 pF/ 1 nF | 3 urad
These are a factor of 10 or more better than our UB based CMCs.
Therefore use AH2700 results, but UB uncertainties (for CMC compliance).

## GRhistory
1999 report | 1999-09-23 | S21341 | E73015 | [5 empty] | 2004 report | 2004-09-10 | S21567 | E7140017 | [5 empty] | 2009 report | 2009-12-01 | S21745 | E36040128 | [5 empty] | 2015 report | 2015-04-30 | S21959 | E96040236
| Dial | Setting | Capacitance | Expanded | k | Loss angle | Expanded | k |  | Dial | Setting | Capacitance | Expanded | k | Loss angle | Expanded | k |  | Dial | Setting | Capacitance | Expanded | k | Loss angle | Expanded | k |  | Dial | Setting | Capacitance | Expanded | k | Loss angle | Expanded | k
|  |  | change | uncertainty |  |  | uncertainty |  |  |  |  | change | uncertainty |  |  | uncertainty |  |  |  |  | change | uncertainty |  |  | uncertainty |  |  |  |  | change | uncertainty |  |  | uncertainty
[6 empty] | (milliradian) | (milliradian) | [7 empty] | (milliradian) | (milliradian) | [7 empty] | (milliradian) | (milliradian) | [7 empty] | (milliradian) | (milliradian)
E1 | 1pF | 10 | 10.2 | 1.2 | 2.4696 | 0.088527 | 0.0346 | 2.1477 |  | 1pF | 10 | 10.6 | 1.2 | 2.4696 | 0.094467 | 0.034599 | 2.1477 |  | 1pF | 10 | 10.3 | 1.2 | 2.4696 | 0.21829 | 0.2037 | 2.1404 |  | 1pF | 10 | 10.3256 | 1.2348 | 2.4695 | 0.092024 | 0.34495 | 2.082
E2 | DIV | 20 | 20.4647 | 1.2348 | 2.4696 | 0.13193 | 0.034559 | 2.1491 |  | DIV | 20 | 20.9974 | 1.2348 | 2.4696 | 0.1762 | 0.034559 | 2.1492 |  | DIV | 20 | 21.0168 | 1.2348 | 2.4696 | 0.29118 | 0.11729 | 2.1438 |  | DIV | 20 | 21.0663 | 1.2348 | 2.4695 | 0.13261 | 0.17817 | 2.0702
E3 |  | 30 | 31.0341 | 1.2348 | 2.4696 | 0.12888 | 0.034533 | 2.1501 |  |  | 30 | 31.4918 | 1.2348 | 2.4696 | 0.16511 | 0.034533 | 2.1501 |  |  | 30 | 31.6982 | 1.2348 | 2.4696 | 0.27445 | 0.089296 | 2.146 |  |  | 30 | 31.7019 | 1.2348 | 2.4694 | 0.052217 | 0.12525 | 2.0637
E4 |  | 40 | 41.5489 | 1.2348 | 2.4696 | 0.15403 | 0.034515 | 2.1508 |  |  | 40 | 41.4958 | 1.2348 | 2.4696 | 0.18314 | 0.034516 | 2.1508 |  |  | 40 | 41.8126 | 1.2348 | 2.4696 | 0.29798 | 0.075987 | 2.1474 |  |  | 40 | 41.8955 | 1.2348 | 2.4694 | 0.049209 | 0.10028 | 2.0612
E5 |  | 50 | 51.4655 | 1.2348 | 2.4695 | 0.023317 | 0.040484 | 2.0755 |  |  | 50 | 51.1505 | 1.2348 | 2.4695 | 0.01955 | 0.040558 | 2.0753 |  |  | 50 | 51.557 | 1.2348 | 2.4695 | 0.031034 | 0.095176 | 2.0672 |  |  | 50 | 51.636 | 1.2348 | 2.4693 | 0.038303 | 0.085959 | 2.0614
E6 |  | 60 | 61.3009 | 1.2348 | 2.4695 | 0.014682 | 0.039422 | 2.079 |  |  | 60 | 60.8898 | 1.2348 | 2.4695 | 0.0082117 | 0.039491 | 2.0786 |  |  | 60 | 61.6953 | 1.2348 | 2.4695 | 0.022693 | 0.083241 | 2.0646 |  |  | 60 | 61.1184 | 1.2348 | 2.4693 | 0.032946 | 0.076623 | 2.0633
E7 |  | 70 | 70.4282 | 1.2348 | 2.4695 | 0.012779 | 0.038662 | 2.0829 |  |  | 70 | 70.2149 | 1.2348 | 2.4695 | 0 | 0.038694 | 2.0827 |  |  | 70 | 70.4245 | 1.2348 | 2.4695 | 0.0142 | 0.075891 | 2.0639 |  |  | 70 | 70.6509 | 1.2348 | 2.4692 | 0.048452 | 0.069923 | 2.0663
E8 |  | 80 | 80.4827 | 1.2348 | 2.4694 | 0.0074552 | 0.038012 | 2.0873 |  |  | 80 | 80.3052 | 1.2348 | 2.4694 | 0 | 0.038026 | 2.0873 |  |  | 80 | 80.2848 | 1.2348 | 2.4694 | 0.0062279 | 0.069657 | 2.0645 |  |  | 80 | 80.2034 | 1.2348 | 2.4691 | 0.04799 | 0.064926 | 2.0699
E9 |  | 90 | 90.4851 | 1.2348 | 2.4694 | 0.0022104 | 0.037628 | 2.0893 |  |  | 90 | 90.0535 | 1.2348 | 2.4694 | 0 | 0.037517 | 2.0917 |  |  | 90 | 90.4232 | 1.2348 | 2.4694 | 0 | 0.064791 | 2.0662 |  |  | 90 | 90.3229 | 1.2348 | 2.4691 | 0.041556 | 0.060881 | 2.0742
E10 |  | '100' | 100.2635 | 1.2348 | 2.4693 | 0 | 0.03708 | 2.0963 |  |  | '100' | 99.8057 | 1.2348 | 2.4693 | -0.01002 | 0.037114 | 2.0958 |  |  | '100' | 100.2685 | 1.2348 | 2.4693 | -0.010971 | 0.061102 | 2.0686 |  |  | '100' | 100.1625 | 1.2348 | 2.469 | 0.025056 | 0.057803 | 2.0785
D1 | 100pF | 1 | 97.49 | 0.016 | 2.1075 | 0.074881 | 0.037187 | 2.0951 |  | 100pF | 1 | 97.145 | 0.016 | 2.1039 | 0.05147 | 0.037211 | 2.0948 |  | 100pF | 1 | 96.977 | 0.015 | 2.1079 | 0.03506 | 0.062242 | 2.0678 |  | 100pF | 1 | 96.96 | 0.02408 | 2.2338 | 0.32282 | 0.058729 | 2.0771
D2 | STEPS | 2 | 195.9764 | 0.030256 | 2.1366 | 0.023983 | 0.035383 | 2.1248 |  | STEPS | 2 | 195.8469 | 0.030359 | 2.1321 | 0.020424 | 0.035387 | 2.1247 |  | STEPS | 2 | 195.8394 | 0.030609 | 2.1249 | -0.0010213 | 0.046049 | 2.0993 |  | STEPS | 2 | 195.9392 | 0.040183 | 2.2424 | 0.047648 | 0.045048 | 2.1123
D3 |  | 3 | 293.3355 | 0.044873 | 2.147 | 0.032046 | 0.034919 | 2.1372 |  |  | 3 | 292.8383 | 0.044978 | 2.1421 | 0.020489 | 0.034923 | 2.1371 |  |  | 3 | 292.6765 | 0.045488 | 2.1304 | 0.0061502 | 0.04163 | 2.1199 |  |  | 3 | 292.7762 | 0.055966 | 2.2449 | 0.13935 | 0.04117 | 2.1293
D4 |  | 4 | 392.6409 | 0.059837 | 2.1517 | 0.016046 | 0.034735 | 2.143 |  |  | 4 | 392.2299 | 0.060013 | 2.1467 | 0.012748 | 0.034738 | 2.1429 |  |  | 4 | 392.2489 | 0.06082 | 2.133 | -0.0086681 | 0.039572 | 2.1315 |  |  | 4 | 392.4936 | 0.072229 | 2.246 | 0.04977 | 0.039311 | 2.1381
D5 |  | 5 | 489.943 | 0.074526 | 2.1542 | 0.02286 | 0.03464 | 2.1462 |  |  | 5 | 489.1993 | 0.07471 | 2.149 | 0.010221 | 0.034651 | 2.1459 |  |  | 5 | 489.044 | 0.075742 | 2.1342 | -0.00593 | 0.038445 | 2.1379 |  |  | 5 | 489.3095 | 0.088022 | 2.2465 | 0.10453 | 0.038276 | 2.1427
D6 |  | 6 | 592.9669 | 0.090096 | 2.1557 | 0.017877 | 0.034587 | 2.1481 |  |  | 6 | 592.3264 | 0.090356 | 2.1505 | 0.0033766 | 0.034701 | 2.1441 |  |  | 6 | 592.0099 | 0.091625 | 2.135 | -0.010642 | 0.037683 | 2.1422 |  |  | 6 | 592.1304 | 0.1048 | 2.2468 | 0.094743 | 0.037568 | 2.1457
D7 |  | 7 | 691.4952 | 0.105 | 2.1566 | 0.023283 | 0.034553 | 2.1494 |  |  | 7 | 691.042 | 0.10534 | 2.1513 | 0.013458 | 0.034558 | 2.1492 |  |  | 7 | 690.8962 | 0.10689 | 2.1355 | -0.0041975 | 0.037181 | 2.1448 |  |  | 7 | 691.1667 | 0.12096 | 2.247 | 0.050106 | 0.037095 | 2.1475
D8 |  | 8 | 788.8593 | 0.11973 | 2.1572 | 0.033594 | 0.03453 | 2.1502 |  |  | 8 | 788.0353 | 0.12008 | 2.1519 | 0.025761 | 0.034531 | 2.1502 |  |  | 8 | 787.7233 | 0.12183 | 2.1358 | 0.0074901 | 0.036819 | 2.1466 |  |  | 8 | 787.9986 | 0.13676 | 2.2471 | 0.083688 | 0.036753 | 2.1487
D9 |  | 9 | 888.1807 | 0.13476 | 2.1576 | 0.043573 | 0.034514 | 2.1508 |  |  | 9 | 887.486 | 0.13519 | 2.1523 | 0.038311 | 0.034514 | 2.1508 |  |  | 9 | 887.3457 | 0.13721 | 2.136 | 0.023216 | 0.036535 | 2.1479 |  |  | 9 | 887.734 | 0.15303 | 2.2472 | 0.050688 | 0.036482 | 2.1496
D10 |  | X | 985.3288 | 0.14947 | 2.1579 | 0.050137 | 0.034503 | 2.1513 |  |  | X | 984.2543 | 0.14989 | 2.1526 | 0.042673 | 0.034503 | 2.1512 |  |  | X | 983.9547 | 0.15213 | 2.1362 | 0.025916 | 0.036317 | 2.1488 |  |  | X | 984.3529 | 0.1688 | 2.2473 | 0.077226 | 0.036273 | 2.1502
C1 | 1000pF | 1 | 0.99505 | 0.00017 | 2.0804 | 0.020099 | 0.040169 | 2.0762 |  | 1000pF | 1 | 0.99471 | 0.00017 | 2.078 | 0.023121 | 0.040172 | 2.0763 |  | 1000pF | 1 | 0.9952 | 0.00016 | 2.0906 | 0.042201 | 0.049981 | 2.0662 |  | 1000pF | 1 | 0.99596 | 0.00017247 | 2.222 | 0.024097 | 0.045354 | 2.0619
C2 | STEPS | 2 | 1.9976 | 0.00031105 | 2.1281 | 0.010011 | 0.03587 | 2.114 |  | STEPS | 2 | 1.9967 | 0.00031216 | 2.1239 | 0.010517 | 0.035847 | 2.1146 |  | STEPS | 2 | 1.9971 | 0.00031376 | 2.1202 | 0.019527 | 0.038708 | 2.0866 |  | STEPS | 2 | 1.9991 | 0.00033526 | 2.2403 | 0.010004 | 0.03751 | 2.1004
C3 |  | 3 | 2.9924 | 0.00045898 | 2.1438 | 0.010025 | 0.035105 | 2.1317 |  |  | 3 | 2.9909 | 0.00046059 | 2.1391 | 0.0080239 | 0.035095 | 2.132 |  |  | 3 | 2.992 | 0.00046576 | 2.1288 | 0.015374 | 0.036484 | 2.1131 |  |  | 3 | 2.9947 | 0.00049746 | 2.2442 | 0.0083478 | 0.035976 | 2.124
C4 |  | 4 | 3.9949 | 0.00060949 | 2.1503 | -0.0025031 | 0.035404 | 2.1236 |  |  | 4 | 3.9927 | 0.00061159 | 2.1453 | 0.0010018 | 0.037183 | 2.0974 |  |  | 4 | 3.9943 | 0.00061976 | 2.1322 | 0.0062586 | 0.035684 | 2.1278 |  |  | 4 | 3.9982 | 0.00066111 | 2.2456 | 0.00075031 | 0.035407 | 2.1354
C5 |  | 5 | 4.9985 | 0.00076076 | 2.1534 | -0.004001 | 0.034876 | 2.1383 |  |  | 5 | 4.9959 | 0.00076339 | 2.1483 | -0.0032025 | 0.03482 | 2.14 |  |  | 5 | 4.9973 | 0.00077423 | 2.1339 | 0.0002001 | 0.035303 | 2.1359 |  |  | 5 | 5.0024 | 0.00082491 | 2.2463 | -0.0047976 | 0.035129 | 2.1413
C6 |  | 6 | 5.9927 | 0.00091086 | 2.1552 | -0.0066745 | 0.034659 | 2.1455 |  |  | 6 | 5.9896 | 0.00091401 | 2.15 | -0.0066779 | 0.034623 | 2.1468 |  |  | 6 | 5.9916 | 0.00092751 | 2.1348 | -0.0011682 | 0.03509 | 2.1407 |  |  | 6 | 5.9974 | 0.00098726 | 2.2467 | -0.0061691 | 0.034971 | 2.1446
C7 |  | 7 | 6.9955 | 0.0010624 | 2.1563 | 0 | 0.03456 | 2.1491 |  |  | 7 | 6.9916 | 0.0010661 | 2.1511 | 0.00057208 | 0.042668 | 2.095 |  |  | 7 | 6.9937 | 0.0010821 | 2.1353 | 0.0035745 | 0.034956 | 2.1438 |  |  | 7 | 7.0007 | 0.001151 | 2.2469 | -0.00028568 | 0.034869 | 2.1468
C8 |  | 8 | 7.9896 | 0.0012128 | 2.157 | 0.0087611 | 0.034633 | 2.1465 |  |  | 8 | 7.9852 | 0.0012169 | 2.1517 | 0.010519 | 0.034582 | 2.1483 |  |  | 8 | 7.9878 | 0.0012355 | 2.1357 | 0.013395 | 0.034867 | 2.1458 |  |  | 8 | 7.9956 | 0.0013133 | 2.2471 | 0.009755 | 0.0348 | 2.1481
C9 |  | 9 | 8.9917 | 0.0013644 | 2.1575 | 0.02113 | 0.034552 | 2.1494 |  |  | 9 | 8.9867 | 0.0013691 | 2.1522 | 0.022254 | 0.034541 | 2.1498 |  |  | 9 | 8.9899 | 0.0013902 | 2.1359 | 0.024248 | 0.034803 | 2.1472 |  |  | 9 | 8.9988 | 0.001477 | 2.2472 | 0.022002 | 0.034751 | 2.149
C10 |  | X | 9.9859 | 0.0015149 | 2.1579 | 0.034047 | 0.034531 | 2.1502 |  |  | X | 9.9803 | 0.00152 | 2.1526 | 0.036069 | 0.034525 | 2.1504 |  |  | X | 9.984 | 0.0015437 | 2.1361 | 0.039461 | 0.034757 | 2.1481 |  |  | X | 9.9938 | 0.0016394 | 2.2472 | 0.036322 | 0.034714 | 2.1497
B1 | .01mF | 1 | 10.01 | 0.009 | 2.1352 | 0.029969 | 0.2564 | 2.2485 |  | 0.01mF | 1 | 10.0058 | 0.001708 | 2.0752 | 0.036977 | 0.040658 | 2.0749 |  | 0.01mF | 1 | 10.0116 | 0.0017 | 2.0885 | 0.040951 | 0.04884 | 2.0674 |  | 0.01µF | 1 | 10.0212 | 0.0016624 | 2.2202 | 0.040912 | 0.044109 | 2.0612
B2 | STEPS | 2 | 20.0017 | 0.0094115 | 2.1107 | 0.039995 | 0.13041 | 2.2152 |  | STEPS | 2 | 19.9948 | 0.0031296 | 2.1229 | 0.026506 | 0.035945 | 2.1125 |  | STEPS | 2 | 20.0005 | 0.0031445 | 2.1196 | 0.029498 | 0.038051 | 2.0844 |  | STEPS | 2 | 20.0207 | 0.0032845 | 2.24 | 0.029469 | 0.036828 | 2.098
B3 |  | 3 | 30.0116 | 0.0098607 | 2.0845 | 0.026655 | 0.089259 | 2.1739 |  |  | 3 | 30.0004 | 0.0046213 | 2.1387 | 0.017999 | 0.035202 | 2.1288 |  |  | 3 | 30.0119 | 0.0046727 | 2.1286 | 0.020991 | 0.036079 | 2.11 |  |  | 3 | 30.0416 | 0.004917 | 2.2441 | 0.020637 | 0.035565 | 2.1208
B4 |  | 4 | 40.0154 | 0.010499 | 2.0652 | 0.042482 | 0.069653 | 2.1356 |  |  | 4 | 40.001 | 0.0061281 | 2.1451 | 0.011749 | 0.035044 | 2.1331 |  |  | 4 | 40.0129 | 0.006209 | 2.1321 | 0.015494 | 0.035475 | 2.1225 |  |  | 4 | 40.053 | 0.0065495 | 2.2456 | 0.015229 | 0.035196 | 2.1298
B5 |  | 5 | 50.0143 | 0.011311 | 2.0553 | 0.033989 | 0.0587 | 2.1059 |  |  | 5 | 49.9958 | 0.00764 | 2.1482 | 0.0086003 | 0.035078 | 2.1323 |  |  | 5 | 50.0236 | 0.0077504 | 2.1338 | 0.011594 | 0.035265 | 2.1276 |  |  | 5 | 50.0693 | 0.0081834 | 2.2463 | 0.010785 | 0.03509 | 2.1326
B6 |  | 6 | 60.0241 | 0.01227 | 2.0531 | 0.028321 | 0.052033 | 2.0855 |  |  | 6 | 60.0013 | 0.0091565 | 2.15 | 0 | 0.034601 | 2.1476 |  |  | 6 | 60.0348 | 0.0092937 | 2.1347 | 0.0054965 | 0.03522 | 2.1288 |  |  | 6 | 60.0901 | 0.0098184 | 2.2467 | 0.0048259 | 0.035101 | 2.1324
B7 |  | 7 | 70.016 | 0.013344 | 2.0561 | 0.032848 | 0.047773 | 2.0731 |  |  | 7 | 69.9899 | 0.010672 | 2.151 | 0.011573 | 0.035239 | 2.1286 |  |  | 7 | 70.0234 | 0.010835 | 2.1353 | 0.015851 | 0.035266 | 2.1278 |  |  | 7 | 70.0893 | 0.01145 | 2.2469 | 0.01498 | 0.035179 | 2.1304
B8 |  | 8 | 80.0248 | 0.014511 | 2.0621 | 0.024991 | 0.044951 | 2.0664 |  |  | 8 | 79.9952 | 0.012191 | 2.1517 | 0.02125 | 0.035327 | 2.1267 |  |  | 8 | 80.0345 | 0.012379 | 2.1357 | 0.025363 | 0.035373 | 2.1254 |  |  | 8 | 80.1099 | 0.013085 | 2.2471 | 0.024341 | 0.035307 | 2.1273
B9 |  | 9 | 90.0277 | 0.015746 | 2.0694 | 0.027768 | 0.043048 | 2.0634 |  |  | 9 | 89.9957 | 0.01371 | 2.1522 | 0.033222 | 0.035482 | 2.1235 |  |  | 9 | 90.0352 | 0.013923 | 2.1359 | 0.03865 | 0.035526 | 2.1223 |  |  | 9 | 90.1209 | 0.014719 | 2.2472 | 0.037948 | 0.035475 | 2.1237
B10 |  | X | 100.0375 | 0.017035 | 2.0771 | 0.039983 | 0.041754 | 2.0626 |  |  | X | 100.0012 | 0.015231 | 2.1526 | 0.048897 | 0.035679 | 2.1198 |  |  | X | 100.0457 | 0.015469 | 2.1361 | 0.054672 | 0.035719 | 2.1186 |  |  | X | 100.1418 | 0.016354 | 2.2472 | 0.053323 | 0.035679 | 2.1198
A1 | 0.1mF | 1 | 0.10003 | 1.7e-05 | 2.0771 | 0.042985 | 0.041755 | 2.0626 |  | 0.1mF | 1 | 0.099994 | 1.7e-05 | 2.0748 | 0.039 | 0.041759 | 2.0626 |  | 0.1mF | 1 | 0.099981 | 1.7e-05 | 2.0882 | 0.039006 | 0.04965 | 2.0592 |  | 0.1µF | 1 | 0.10003 | 1.6523e-05 | 2.2199 | 0.042985 | 0.045009 | 2.052
A2 | STEPS | 2 | 0.20038 | 3.1239e-05 | 2.1269 | 0.044914 | 0.040858 | 2.073 |  | STEPS | 2 | 0.20036 | 3.1361e-05 | 2.1229 | 0.041922 | 0.040858 | 2.073 |  | STEPS | 2 | 0.20036 | 3.1502e-05 | 2.1196 | 0.046414 | 0.042709 | 2.0568 |  | STEPS | 2 | 0.20049 | 3.2817e-05 | 2.24 | 0.051374 | 0.04162 | 2.0647
A3 |  | 3 | 0.3004 | 4.6091e-05 | 2.1434 | 0.036616 | 0.046827 | 2.1029 |  |  | 3 | 0.30035 | 4.6268e-05 | 2.1387 | 0.036622 | 0.046823 | 2.1028 |  |  | 3 | 0.30034 | 4.6763e-05 | 2.1286 | 0.046611 | 0.047486 | 2.0928 |  |  | 3 | 0.30052 | 4.9114e-05 | 2.2441 | 0.05091 | 0.04711 | 2.0985
A4 |  | 4 | 0.40073 | 6.1147e-05 | 2.1501 | 0.039926 | 0.055546 | 2.1532 |  |  | 4 | 0.4007 | 6.1386e-05 | 2.1451 | 0.046666 | 0.055542 | 2.1532 |  |  | 4 | 0.40072 | 6.2181e-05 | 2.1321 | 0.058642 | 0.055824 | 2.1469 |  |  | 4 | 0.40097 | 6.5493e-05 | 2.2456 | 0.060602 | 0.055685 | 2.1506
A5 |  | 5 | 0.50055 | 7.6188e-05 | 2.1533 | 0.059932 | 0.065718 | 2.2059 |  |  | 5 | 0.50049 | 7.6482e-05 | 2.1482 | 0.067331 | 0.065711 | 2.2059 |  |  | 5 | 0.50055 | 7.7553e-05 | 2.1338 | 0.075913 | 0.065848 | 2.2018 |  |  | 5 | 0.50087 | 8.179e-05 | 2.2463 | 0.079659 | 0.065807 | 2.2044
A6 |  | 6 | 0.60058 | 9.1289e-05 | 2.1551 | 0.063269 | 0.07676 | 2.2524 |  |  | 6 | 0.60049 | 9.1637e-05 | 2.15 | 0.059948 | 0.076749 | 2.2524 |  |  | 6 | 0.60053 | 9.2966e-05 | 2.1347 | 0.073598 | 0.076818 | 2.2497 |  |  | 6 | 0.60091 | 9.8112e-05 | 2.2467 | 0.084869 | 0.076824 | 2.2514
A7 |  | 7 | 0.70093 | 0.00010645 | 2.1563 | 0.068478 | 0.088329 | 2.2906 |  |  | 7 | 0.70085 | 0.00010687 | 2.151 | 0.072765 | 0.08832 | 2.2906 |  |  | 7 | 0.70092 | 0.00010845 | 2.1353 | 0.089307 | 0.088359 | 2.2888 |  |  | 7 | 0.70136 | 0.0001145 | 2.2469 | 0.10123 | 0.088393 | 2.29
A8 |  | 8 | 0.80096 | 0.00012158 | 2.157 | 0.079901 | 0.10015 | 2.321 |  |  | 8 | 0.80085 | 0.00012205 | 2.1517 | 0.086529 | 0.10014 | 2.3209 |  |  | 8 | 0.8009 | 0.00012388 | 2.1357 | 0.10712 | 0.10016 | 2.3198 |  |  | 8 | 0.80139 | 0.00013083 | 2.2471 | 0.12079 | 0.10021 | 2.3206
A9 |  | 9 | 0.90129 | 0.00013676 | 2.1575 | 0.092087 | 0.11219 | 2.3451 |  |  | 9 | 0.90119 | 0.00013729 | 2.1522 | 0.10796 | 0.11218 | 2.3451 |  |  | 9 | 0.90128 | 0.00013938 | 2.1359 | 0.13059 | 0.11219 | 2.3443 |  |  | 9 | 0.90184 | 0.00014722 | 2.2472 | 0.1437 | 0.11226 | 2.3449
A10 |  | X | 1.0013 | 0.0001519 | 2.1578 | 0.11384 | 0.1243 | 2.3643 |  |  | X | 1.0012 | 0.00015248 | 2.1526 | 0.12884 | 0.12429 | 2.3642 |  |  | X | 1.0013 | 0.00015481 | 2.1361 | 0.149 | 0.12429 | 2.3636 |  |  | X | 1.0019 | 0.00016354 | 2.2472 | 0.16519 | 0.12437 | 2.3641
| Dial | Setting | Capacitance | Expanded | k | Loss angle | Expanded | k |  | Dial | Setting | Capacitance | Expanded | k | Loss angle | Expanded | k |  | Dial | Setting | Capacitance | Expanded | k | Loss angle | Expanded | k |  | Dial | Setting | Capacitance | Expanded | k | Loss angle | Expanded | k
|  |  |  | uncertainty |  |  | uncertainty | [5 empty] | uncertainty |  |  | uncertainty | [5 empty] | uncertainty |  |  | uncertainty | [5 empty] | uncertainty |  |  | uncertainty
[6 empty] | (milliradian) | (milliradian) | [7 empty] | (milliradian) | (milliradian) | [7 empty] | (milliradian) | (milliradian) | [7 empty] | (milliradian) | (milliradian)
E0 | All dials | 0 | 49.4 | 1.5 | 2.4696 | 0 | 0.034666 | 2.1453 |  | All dials | 0 | 49.1 | 1.5 | 2.4696 | 0 | 0.034669 | 2.1453 |  | All dials | 0 | 49.1 | 1.5 | 2.4696 | 3.2825 | 0.035596 | 2.1352 |  | All dials | 0 | 48.9946 | 1.4818 | 2.4696 | 2.9562 | 0.069554 | 2.0988
A10
| C | uC | D | uD
1999-09-23 | 1.0013 | 0.0001519 | 0.11384 | 0.1243
2004-09-10 | 1.0012 | 0.00015248 | 0.12884 | 0.12429
2009-12-01 | 1.0013 | 0.00015481 | 0.149 | 0.12429
2015-04-30 | 1.0019 | 0.00016354 | 0.16519 | 0.12437